SLAC Magnetic Measurement Plan for

LCLS-II Quad Corrector CSQ02B (DSG-000050058)

This traveler is intended to cover mechanical fiducialization and magnetic measurements of the new quadrupole corrector CSQ02B magnet, which is needed to replace the flexible printed circuit board skew quad of the same MAD name.  It will be designated a 3.37Q1.2 type magnet.  The magnet will be installed in the LCLS-II electron gun area.  The quadrupole corrector CSQ02B has a positive polarity.

Receiving:

The following information is to be noted upon receipt of the magnet by the SLAC MM group:

| Received by (initials): | SDA |
| --- | --- |
| Date received (mm/dd/yyyy): | 5/11/2026 |
| SLAC barcode number: | 4271 |
| Serial number on the magnet: | None |

Preparation:

A beam direction arrow, with text “beam direction”, is to be applied to the magnet with a sticker

| Beam-direction arrow in place (initials): | SDA |
| --- | --- |

Fiducialization:

There is no fiducialization for this magnet.  Alignment is to be made using a level.

Magnetic Measurements:

Enter URL of on-line magnetic measurements data (please modify or correct if necessary):

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/LCLS-II/Quad/4271

- Measure the inductance and resistance of the magnet:

| Inductance of coil (mH): | 58.2 µH @ 100 Hz |
| --- | --- |
| Resistance of coil (Ohms): | 0.0506 Ohm |
| Ambient temperature in degrees C | 23.4 oC |

- Determine the connection polarity (with main supply outputting positive current) which produces a “positive” field polarity for CSQ02B as shown below:

Figure 1.  The magnet CSQ02B is “positive”.

- Mark the polarity near the magnet leads with clear “+” and “” labels as shown above.

| Magnet polarity chosen from Fig. 1 is (P): | P |
| --- | --- |

- Run the magnet up to 6 A for 2 hours to warm it up (magnet coil temperature).

| Top Coil (°C) | 29.3 °C |
| --- | --- |
| Left Coil (°C) | 27.6 °C |
| Right Coil (°C) | 28.4 °C |
| Bottom Coil (°C) | 26.4 °C |
| Ambient temperature (°C): | 23.8 °C |

- Standardize the magnet, starting from zero go to +6 A, then go through 1 full cycle from +6 A to -6 A, finally ending at -6 A, with a flat-top pause time (at both -6 A and +6 A) of 10 seconds.  Use a linear ramp rate of 1 A/sec. Record the ramp rate used.

| Standardization complete (initials): | SDA |
| --- | --- |
| Ramp type and rate used (A/sec): | Linear 1 A/sec |

- Measure the integral gradient, Gdl, from -6 to 6 A and then from 6 A to -6 A in 1 amp steps. Measure harmonics at +/-6 A.

| Filename of Gdl data: | Strdat.ru1 |
| --- | --- |
| Filename of Harmonics data: | Hardat.ru1 |

- Measure the Gdl field at with the power supply off.

| Filename of Gdl field at with the power supply off | Strdat.ru2 |
| --- | --- |

- Confirm the pole-tip field using a Hall probe at the surface of the inside clamping ring, at an excitation current of 6 A.

| Hall probe pole-tip field at 6 A (mean of 4 poles): | 29.1 +/- 0.5 G @ 6.00298 |
| --- | --- |

- Upon completion of tests, email URL of on-line data to Mark Woodley. Mark Woodley will determine if the magnet is accepted.  Upon acceptance of magnet, analysis data will be placed in on-line data folder.

| Magnet accepted and Analysis file(s) put into on-line data folder (initials): | SDA |
| --- | --- |
| Assigned beamline location (MAD-deck name): | CSQ02B |