SLAC Magnetic Measurement Plan and Traveler for LCLS-II Injector Laser Heater Undulator, (May 4, 2018)

This traveler covers magnetic measurements of the LCLS laser-heater permanent magnet planar undulator needed for LCLS-II.  The undulator magnet is about 0.5 m long with 9 periods, each with a 5.4-cm length.

Receiving:

The following information is to be noted upon receipt of the magnets by the SLAC MFD group:

| Received by (initials): | YL |
| --- | --- |
| Date received (dd-mmm-yyyy): | 05/12/2021 |
| SLAC barcode number: | N/A |
| Serial number on the magnet (if available): | LHU-0002 |

| Place a barcode sticker on the magnet and also duplicate the barcode sticker here | N/A |
| --- | --- |

Preparation:

A beam direction arrow, with text “beam direction”, is to be applied to the top and/or connector side of the magnet with a sticker supplied by LCLS (H. Alvarez will determine the direction).

| Beam-direction arrow in place (initials): | YL |
| --- | --- |

Magnetic Measurements:

Enter URL of on-line magnetic measurements data (please modify or correct if necessary):

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/LCLS-II/Undulator/LHU-0002/DATASET0001/Final Results/Zscans/....

- Allow the magnet to come to an equilibrium temperature.

- Install and calibrate the LVDT and possibly the gap motor control so that the readings from the same LVDT that is used in the final installation can be recorded along with the K values and field integrals.

- At a gap setting corresponding to an undulator K of 0.90 (B = 1.8 kG), determine the magnetic undulator axis and fiducialize the magnet to external tooling balls so that the undulator can be aligned to the beam coordinates along its central axis to an accuracy of <0.5 mm horizontally and vertically.

- Measure the undulator K value and the precise gap settings at the gap height settings associated with the following K values (0.8, 0.9 1.0, 1.2, 1.4, 1.6, and fully open at about 100 mm).  The gap at K = 0.90 should be about 41 mm.  Please record the gap LVDT reading which is associated with each gap setting, after calibrating the LVDT with appropriate polynomial.

| Configuration | Measured K Km | Measured gap (mm) | LVDT gap reading (mm) |
| --- | --- | --- | --- |
| K 0.80 | 0.800566 | 42.26 | 43.0922 |
| K 0.90 | 0.900115 | 40.30 | 41.1638 |
| K 1.0 | 0.997059 | 38.58 | 39.5132 |
| K 1.2 | 1.199481 | 35.66 | 36.5047 |
| K 1.4 | 1.400030 | 33.14 | 34.0174 |
| K 1.6 | 1.578202 | 31.19 | 32.0892 |
| K max. (i.e., min. gap) | 2.493565 | 23.33 | 24.6147 |
| Directory name: | …./On Axis/…. |  |  |

Measurement temperature Tm= 19.97 C°. The K in the tunnel Kt = Km(1-10-3(T-Tm)).

- Measure the 1st and 2nd field integrals all along the undulator length for both the horizontal and vertical plane in the following conditions in the mid-plane. Directory name: …./Good Field Region/ …

| Full gap height set for K = 0.90 (x = 15 mm): | Filename: …/Run 001/007gap041.172x-15.00y+00.00 |
| --- | --- |
| Full gap height set for K = 0.90 (x = 10 mm): | Filename: …/Run 001/006gap041.172x-10.00y+00.00 |
| Full gap height set for K = 0.90 (x = 5 mm): | Filename: …/Run 001/005gap041.172x-05.00y+00.00 |
| Full gap height set for K = 0.90 (x = 0): | Filename: …/Run 001/004gap041.172x+00.00y+00.00 |
| Full gap height set for K = 0.90 (x = 5 mm): | Filename: …/Run 001/003gap041.172x+05.00y+00.00 |
| Full gap height set for K = 0.90 (x = 10 mm): | Filename: …/Run 001/002gap041.172x+10.00y+00.00 |
| Full gap height set for K = 0.90 (x = 15 mm): | Filename: …/Run 001/001gap041.172x+15.00y+00.00 |
| Full gap height set for K = 0.800 (x =  mm): | Filename: …/Run 001/011gap043.102x+00.00y+00.00 |
| Full gap height set for K = 1.600 (x =  mm): | Filename: …/Run 001/009gap031.875x+00.00y+00.00 |
| Full gap height set for K = max. (x =  mm): | Filename: …/Run 001/008gap025.000x+00.00y+00.00 |

- Upon completion of tests, email URL of on-line data to Mark Woodley. Mark Woodley will determine if the magnet is accepted.  Upon acceptance of magnet, analysis data will be placed in on-line data folder.

| Magnet accepted and Analysis file(s) put into on-line data folder (initials): |  |
| --- | --- |
| Assigned beamline location (MAD-deck name): |  |