Angles between Hall probes

Roll

YHi guys,

Soldering would be fine if you keep the soldering iron on as long as you really need. Do not leave it on for all day.

Regards,

Yurii

XHi guys,

Soldering would be fine if you keep the soldering iron on as long as you really need. Do not leave it on for all day.

Regards,

Yurii

ZHi guys,

Soldering would be fine if you keep the soldering iron on as long as you really need. Do not leave it on for all day.

Regards,

Yurii

+α

|  | Probe 1 (S/N 120) |  | Probe 2 (S/N 119) |  |
| --- | --- | --- | --- | --- |
|  | X | Y | X | Y |
| Stage readings() | 306.61 | 214.81 | 306.07 | 214.36 |
|  | 306.62 | 214.82 | 306.07 | 214.36 |
|  |  | 214.82 | 306.07 |  |
| Average() | 306.62 | 214.82 | 306.07 | 214.36 |
| Differences() | 0 | 91.80(+1.80) | +0.55 | 92.26(+2.26) |
|  | -1.80 | 0 | -1.25 | +0.46 |

XHi guys,

Soldering would be fine if you keep the soldering iron on as long as you really need. Do not leave it on for all day.

Regards,

Yurii

ZHi guys,

Soldering would be fine if you keep the soldering iron on as long as you really need. Do not leave it on for all day.

Regards,

Yurii

YHi guys,

Soldering would be fine if you keep the soldering iron on as long as you really need. Do not leave it on for all day.

Regards,

Yurii

Pitch

+φ

|  | V ( By) | V ( By) | Zero offset | ΔV | Angle (rad) | Angle() |
| --- | --- | --- | --- | --- | --- | --- |
| Probe 1 | -11.56 mV | +12.91mV | 0.68mV | 0 | -0.01055 | -0.60 |
| Probe 2 | -23.97 mV | +24.37mV | 0.20mV | -11.93mV | -0.02080 | -1.19 |