## Slide 1: Tube Laser Alignment System
Georg Gassner

## Slide 2: Delta Undulator Project
Period 32 mm, Length 3.2m
Gap (bore) 6.4mm
PM material NdFeB, grade N40UH
Vacuum
Chamber ID 5.0mm

## Slide 3: Alignment Scope
Determine the x- and y- location of a Hall probe assembly at different z-locations inside the vacuum tube allowing simultaneous position and magnetic measurement.

## Slide 4: Concept: x-Position
CCD Camera
Retroreflector
Laser
Beam Splitter
a)
CCD Camera
Laser
Beam Splitter
Retroreflector moved by Δx
2Δx
2Δx
b)
Δx

## Slide 5: Concept: x-y-Position
Incoming beam: Δx
Δy
Outgoing beam: -Δx
-Δy

## Slide 6: Concept: z-Position
We already use a laser and a retroreflector => replace laser head with interferometer laser and we get distance.
CCD Camera
Retroreflector
Laser Head
Interferometer
Beam Splitter
Receiver

## Slide 7: Retroreflector (1)
Prism cut from solid
1.5” CCR with
diamond core drill

## Slide 8: Retroreflector (2)
- Prism Ø 4.5mm
- Prism and Hall Probe might not be
centered => roll stability important
(carbon fiber tube pusher system)

## Slide 9: Laser - Hardware
Agilent 5517B Laser Head
  Helium-Neon (632nm), Continuous wave
  Beam diameter 6mm
  Beam divergence ~0.15mrad

## Slide 10: Laser - Beam Size Reduction
d1=6mm
d2=2mm
1=0.15mrad
2=0.45mrad
f1=f2

## Slide 11: Laser - Beam Focusing
To move the waist of the laser beam out, add focusing lens or move one of the lenses
3
waist

## Slide 12: Laser Positional Stability

## Slide 13: CCD Image
Point Grey - FL2-20S4M-C
  1624x1224 pixel (~7x5mm)
  4.4µm pixel size
Natural Density Filter (OD=2) to protect CCD

## Slide 14: CCD Images

## Slide 15: Image Processing (1)
Cross Correlation of two images =>DX, DY
Manuel Guizar-Sicairos, Samuel T. Thurman, and James R. Fienup, "Efficient subpixel image registration algorithms," Opt. Lett. 33, 156-158 (2008)

## Slide 16: Image Processing (2)
DY=0.4 pixel
DX=-0.2 pixel

## Slide 17: Measurement System (1)
CCD Camera
3.2m vacuum tube
Retroreflector
Laser Head
Mirror
Interferometer
Beam Splitter
Receiver
Focusing
Lenses
Ø4.8mm
Mirror
Hall Probe

## Slide 18: Measurement System (2)

## Slide 19: Results – Retroreflector Position
Move the retroreflector from 0-3000mm through the vacuum chamber and back

## Slide 20: Summary
First tests indicate that sub 50µm accuracy is achievable.
Can be triggered together with magnetic measurements.
Effects of atmosphere need further investigation especially when system is moving.