## Slide 1: DELTA Quadrant Tuning
Y. Levashov

## Slide 2: What to tune?
Y
Z
X
Y
Z
X
Beam axis:

## Slide 3: Fields from a single magnet array
and  200, < 0.5µm
Y
X
- a single array could be treated as an undulator.
No access to the magnets when assembled.
DELTA is PPM device:
Pros. – superposition works, field at each point is a sum of fields from all magnets. Quadrants could be tuned separately.
Cons. – no iron shims could be used for tuning, virtual shimming.
Challenge for tuning – probe positioning (setting K).

## Slide 4: Field change from magnet motion
Bx field change from single block motion in X
By field change from single block motion in Y
Shim 0.025mm Y block → 12 µTm
0.025mm Y block shift in x → 4 µTm
0.125mm shift of a pair of Y blocks in y → 13 degrees of phase change
0.125mm shift of Z block in y → 4 degrees of phase change
How to tune:
- Move Y-magnets at X-trajectory kicks’ locations up/down to correct X-trajectory.
- Move Y-magnets at Y-trajectory kicks’ locations left/right to correct Y-trajectory.
- Move Y-magnets in pairs or Z-magnets up/down to correct phase errors and K.

## Slide 5: Quadrant tuning process
Find magnetic axis location, place probe on the axis.
Move probe along the axis. Trigger DMMs at z interval.
Apply zero drift and offset corrections to the voltages.
Apply calibrations. Calculate fields.
Calculate trajectories, phase errors and K value
Tune trajectories
Tune phase errors and K; make K the same for all four quadrants
Check magnet mechanical positions on CMM
Iterate

## Slide 6: Senis 2-D probe
Ceramic pad
Plastic cover

## Slide 7: Alignment to the bench (X & yaw)
Chart: By vs probe X position
Do Hall probe scan in X at each pole location
Truncate data ± 3mm from center
Fit parabola, find center
Fit a straight line through all pole centers, find x and yaw, correct yaw
if necessary.

## Slide 8: Alignment to the bench (Y)
Find center of fiducialization magnet with the Hall probe
Measure difference in height between the magnet and the reference block by optical tools.
Measure difference in height between the reference block and the quadrant by Keyence.
Set quadrant by using cam movers.
ΔY = 30μm  ΔB/B  0.5%, relative error quadrant-to-quadrant  0.1%

## Slide 9
X-blocks flipped 180 degrees
Check for assembly errors
Correct

## Slide 10
X-trajectory tuning
Before
After

## Slide 11
Y-trajectory tuning
Before
After

## Slide 12
Phase tuning
Move block up
Move block down

## Slide 13: 3.2m DELTA tuning summary
Quadrants were not straight enough during the tuning, which resulted in big phase errors. Max bow of 18µm and max pitch of 8µm allowed to keep phase errors below 5 degrees. Straightening the quadrant strongback on the tuning bench is required.
Since two of the quadrants flipped during assembly, more attention to be paid to first integral tuning – to be set to zero as close as possible. Otherwise the first integral will result in big second integral for flipped quadrants.
Three iterations per quadrant required to complete the tuning. One iteration = 1 day, which adds to 3 weeks needed for tuning all the quadrants.
Number of field scans/iteration average:
Q1 – 11 (initial tuning, learning curve)
Q2 – 15 (more time spent to tune phase errors)
Q3 - 9
Q4 – 8
Wire measurements desirable to apply corrections for planar Hall effect.