{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"not_rendered":{"charts":0,"images":0,"macros":1,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":197,"n_cols":8,"rows_kept":197,"cut":false,"hidden":false,"merged":[]}
["MMF",null,null,"PM Checks"]
[]
["\"X\""]
[]
["Undulator","SN 13"]
["Date:","2015-07-02"]
["Crew:","FG, BR"]
[]
["BUCK IN "]
["Far Value",0,"PM1 X1"]
["Near Value",0,"PM 4 X1"]
["1st Far",15]
["1st. Near",16.227]
["2nd. Far",16.613]
["2nd Near",16.821000000000002]
["Far Should Read",{"$f":"BuckIn(B10,B11,B12,B13,B14,B15)","v":16.942248282630032}]
["Near Should Read",16.942248282630032]
["LOS",16.942248282630032]
[]
["Undulator "]
[]
[]
["1)","2)","3)","4)"]
[21.001000000000001,20.994,20.998000000000001,21.001000000000001]
[1,1,1,1]
[{"$f":"SUM(A24:A25)","v":22.001000000000001},{"$f":"SUM(B24:B25)","v":21.994},{"$f":"SUM(C24:C25)","v":21.998000000000001},{"$f":"SUM(D24:D25)","v":22.001000000000001}]
[{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032}]
[{"$f":"SUM(A27-A26)","v":-5.0587517173699688},{"$f":"SUM(B27-B26)","v":-5.0517517173699673},{"$f":"SUM(C27-C26)","v":-5.0557517173699686},{"$f":"SUM(D27-D26)","v":-5.0587517173699688}]
[]
["5)","6)","7)","8)"]
[11.223000000000001,11.225,11.224,11.223000000000001]
[1,1,1,1]
[{"$f":"SUM(A31:A32)","v":12.223000000000001},{"$f":"SUM(B31:B32)","v":12.225},{"$f":"SUM(C31:C32)","v":12.224},{"$f":"SUM(D31:D32)","v":12.223000000000001}]
[{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032}]
[{"$f":"SUM(A34-A33)","v":4.7192482826300317},{"$f":"SUM(B34-B33)","v":4.7172482826300328},{"$f":"SUM(C34-C33)","v":4.7182482826300323},{"$f":"SUM(D34-D33)","v":4.7192482826300317}]
[]
[]
[null,"Reference","Block"]
[]
["A)","B)","C)","D)"]
[22.555,18.024000000000001,16.265000000000001,16.265000000000001]
[1,1,1,1]
[{"$f":"SUM(A41:A42)","v":23.555},{"$f":"SUM(B41:B42)","v":19.024000000000001},{"$f":"SUM(C41:C42)","v":17.265000000000001},{"$f":"SUM(D41:D42)","v":17.265000000000001}]
[{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032}]
[{"$f":"SUM(A44-A43)","v":-6.6127517173699673},{"$f":"SUM(B44-B43)","v":-2.0817517173699684},{"$f":"SUM(C44-C43)","v":-0.32275171736996811},{"$f":"SUM(D44-D43)","v":-0.32275171736996811}]
[]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"X\""]
[]
[{"$f":"B18","v":16.942248282630032},"LOS"]
["Undulator",{"$f":"B5","v":"SN 13"}]
["Date:",{"$f":"B6","v":"2015-07-02"}]
["Crew:",{"$f":"B7","v":"FG, BR"}]
[]
[]
[null,"PM4"]
[]
["X1) Buck In","Y1)","Y2)"]
[null,19.542000000000002,19.541]
[null,1,1]
[null,{"$f":"SUM(B64:B65)","v":20.542000000000002},{"$f":"SUM(C64:C65)","v":20.541}]
[null,{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032}]
[0,{"$f":"SUM(B67-B66)","v":-3.5997517173699691},{"$f":"SUM(C67-C66)","v":-3.5987517173699679}]
[]
[]
[null,"PM3"]
[]
["X1)","Y1)","Y2)"]
[17.408000000000001,21.001999999999999,21]
[1,1,1]
[{"$f":"SUM(A74:A75)","v":18.408000000000001},{"$f":"SUM(B74:B75)","v":22.001999999999999},{"$f":"SUM(C74:C75)","v":22}]
[{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032}]
[{"$f":"SUM(A77-A76)","v":-1.4657517173699688},{"$f":"SUM(B77-B76)","v":-5.0597517173699664},{"$f":"SUM(C77-C76)","v":-5.0577517173699675}]
[]
[]
[null,"PM 2"]
["X1)","Y1)","Y2)"]
[17.396999999999998,20.99,20.994]
[1,1,1]
[{"$f":"SUM(A83:A84)","v":18.396999999999998},{"$f":"SUM(B83:B84)","v":21.99},{"$f":"SUM(C83:C84)","v":21.994}]
[{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032}]
[{"$f":"SUM(A86-A85)","v":-1.454751717369966},{"$f":"SUM(B86-B85)","v":-5.047751717369966},{"$f":"SUM(C86-C85)","v":-5.0517517173699673}]
[]
[]
[null,"PM1"]
["X1) Buck In","Y1)","Y2)"]
[null,19.541,19.535]
[null,1,1]
[null,{"$f":"SUM(B92:B93)","v":20.541},{"$f":"SUM(C92:C93)","v":20.535}]
[null,{"$f":"B18","v":16.942248282630032},{"$f":"B18","v":16.942248282630032}]
[0,{"$f":"SUM(B95-B94)","v":-3.5987517173699679},{"$f":"SUM(C95-C94)","v":-3.5927517173699677}]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"Y\"",null,"PM4Y1","PM1Y1"]
[null,null,9.1479999999999997,9.1579999999999995]
[null,null,1,1]
["Undulator",{"$f":"B5","v":"SN 13"},{"$f":"SUM(C104+C105)","v":10.148},{"$f":"SUM(D104+D105)","v":10.157999999999999}]
["Date:",{"$f":"B6","v":"2015-07-02"}]
["Crew:",{"$f":"B7","v":"FG, BR"}]
[null,null,"HI =",{"$f":"AVERAGE(C106,D106)","v":10.152999999999999}]
[]
["PM4",null,"PM3"]
[]
["Y1)","X1)","Y2)","X1"]
[{"$f":"C104","v":9.1479999999999997},12.75,9.1590000000000007,12.762]
[1,1,1,1]
[{"$f":"SUM(A114:A115)","v":10.148},{"$f":"SUM(B114:B115)","v":13.75},{"$f":"SUM(C114:C115)","v":10.159000000000001},{"$f":"SUM(D114:D115)","v":13.762}]
[{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999}]
[{"$f":"SUM(A117-A116)","v":4.9999999999990052E-3},{"$f":"SUM(B117-B116)","v":-3.5970000000000013},{"$f":"SUM(C117-C116)","v":-6.0000000000020037E-3},{"$f":"SUM(D117-D116)","v":-3.6090000000000018}]
[]
[]
["PM2",null,"PM1"]
[]
["Y1)","X1)","Y1)","X1)"]
[9.16,12.762,{"$f":"D104","v":9.1579999999999995},12.759]
[1,1,1,1]
[{"$f":"SUM(A124:A125)","v":10.16},{"$f":"SUM(B124:B125)","v":13.762},{"$f":"SUM(C124:C125)","v":10.157999999999999},{"$f":"SUM(D124:D125)","v":13.759}]
[{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999}]
[{"$f":"SUM(A127-A126)","v":-7.0000000000014495E-3},{"$f":"SUM(B127-B126)","v":-3.6090000000000018},{"$f":"SUM(C127-C126)","v":-5.0000000000007816E-3},{"$f":"SUM(D127-D126)","v":-3.6060000000000016}]
[]
[]
[]
[]
[]
[]
["Top Of Gage Block",null,"Probe Location in (X)",null,null,null,"Off set to probe tip",-5.4640000000000004]
[null,null,null,null,null,null,"Blade Reading",28.001999999999999]
[14.286,null,-5.4640000000000004,"LOS",null,null,"LOS",16.942]
[{"$f":"D109","v":10.152999999999999},null,-0.05,"Mic To Plus Edge"]
[{"$f":"SUM(A138-A137)","v":-4.1330000000000009},null,{"$f":"SUM(C137+C138)","v":-5.5140000000000002},null,null,null,"Blade to LOS",{"$f":"H136+H137","v":44.944000000000003}]
[null,null,-3.9E-2,null,null,null,"LOS to probe tip",{"$f":"H139-H135","v":50.408000000000001}]
[null,null,{"$f":"(C139+C140)","v":-5.5529999999999999},"Probe Location"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"Y\""]
[]
["Undulator",{"$f":"B5","v":"SN 13"}]
["Date:",{"$f":"B107","v":"2015-07-02"}]
["Crew:",{"$f":"B7","v":"FG, BR"}]
[]
[]
["HI =",{"$f":"D109","v":10.152999999999999}]
[]
["Undulator"]
[]
["1)","2)","3)","4)"]
[5.6710000000000003,5.6710000000000003,5.6769999999999996,5.6749999999999998]
[1,1,1,1]
[{"$f":"SUM(A165:A166)","v":6.6710000000000003},{"$f":"SUM(B165:B166)","v":6.6710000000000003},{"$f":"SUM(C165:C166)","v":6.6769999999999996},{"$f":"SUM(D165:D166)","v":6.6749999999999998}]
[{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999}]
[{"$f":"SUM(A168-A167)","v":3.4819999999999984},{"$f":"SUM(B168-B167)","v":3.4819999999999984},{"$f":"SUM(C168-C167)","v":3.4759999999999991},{"$f":"SUM(D168-D167)","v":3.4779999999999989}]
[]
[]
["5)","6)","7)","8)"]
[12.763,12.768000000000001,12.766999999999999,12.772]
[1,1,1,1]
[{"$f":"SUM(A173+A174)","v":13.763},{"$f":"SUM(B173+B174)","v":13.768000000000001},{"$f":"SUM(C173+C174)","v":13.766999999999999},{"$f":"SUM(D173+D174)","v":13.772}]
[{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999}]
[{"$f":"SUM(A176-A175)","v":-3.6100000000000012},{"$f":"SUM(B176-B175)","v":-3.615000000000002},{"$f":"SUM(C176-C175)","v":-3.6140000000000008},{"$f":"SUM(D176-D175)","v":-3.6190000000000015}]
[]
[]
[]
["Reference Block"]
[]
["A)","B)","C)","D)"]
[9.3970000000000002,null,11.54,14.621]
[1,null,1,1]
[{"$f":"SUM(A184+A185)","v":10.397},null,{"$f":"SUM(C184+C185)","v":12.54},{"$f":"SUM(D184+D185)","v":15.621}]
[{"$f":"D109","v":10.152999999999999},null,{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999}]
[{"$f":"SUM(A187-A186)","v":-0.24400000000000155},0,{"$f":"SUM(C187-C186)","v":-2.3870000000000005},{"$f":"SUM(D187-D186)","v":-5.4680000000000017}]
[]
[]
["Flats"]
[]
["U/S",null,"D/S"]
["(+)","(-)","(+)","(-)"]
[8.4429999999999996,8.4420000000000002,8.4450000000000003,8.4440000000000008]
[{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999},{"$f":"D109","v":10.152999999999999}]
[{"$f":"SUM(A196-A195)","v":1.7099999999999991},{"$f":"SUM(B196-B195)","v":1.7109999999999985},{"$f":"SUM(C196-C195)","v":1.7079999999999984},{"$f":"SUM(D196-D195)","v":1.7089999999999979}]

{"sheet":"Sheet2","index":1,"n_rows":42,"n_cols":8,"rows_kept":42,"cut":false,"hidden":false,"merged":[]}
[]
["Undulator"," ",{"$f":"Sheet1!B5","v":"SN 13"}]
["Date"," ",{"$f":"Sheet1!B6","v":"2015-07-02"}]
["Crew"," ",{"$f":"Sheet1!B7","v":"FG, BR"}]
[]
[]
["Undulator Values"]
[]
["TB","Y (in)","Y (mm)",null,null,"TB","X (in)","X (mm)"]
[1,{"$f":"Sheet1!A169","v":3.4819999999999984},{"$f":"B10*25.4","v":88.442799999999949},null,null,5,{"$f":"Sheet1!A35","v":4.7192482826300317},{"$f":"G10*25.4","v":119.86890637880281}]
[2,{"$f":"Sheet1!B169","v":3.4819999999999984},{"$f":"B11*25.4","v":88.442799999999949},null,null,6,{"$f":"Sheet1!B35","v":4.7172482826300328},{"$f":"G11*25.4","v":119.81810637880282}]
[3,{"$f":"Sheet1!C169","v":3.4759999999999991},{"$f":"B12*25.4","v":88.290399999999977},null,null,7,{"$f":"Sheet1!C35","v":4.7182482826300323},{"$f":"G12*25.4","v":119.84350637880281}]
[4,{"$f":"Sheet1!D169","v":3.4779999999999989},{"$f":"B13*25.4","v":88.341199999999972},null,null,8,{"$f":"Sheet1!D35","v":4.7192482826300317},{"$f":"G13*25.4","v":119.86890637880281}]
[]
[]
["PM Values"]
[]
[null,null,"Upstream"," ","(PM 3)",null,"Checks"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"X (mm)","Y (mm)"]
["Y2",{"$f":"Sheet1!C78","v":-5.0577517173699675},{"$f":"Sheet1!C118","v":-6.0000000000020037E-3},{"$f":"B20*25.4","v":-128.46689362119716},{"$f":"C20*25.4","v":-0.15240000000005088},null,{"$f":"H10-D21","v":157.09900000000002},{"$f":"C10-E20","v":88.595200000000006}]
["X1",{"$f":"Sheet1!A78","v":-1.4657517173699688},{"$f":"Sheet1!D118","v":-3.6090000000000018},{"$f":"B21*25.4","v":-37.230093621197206},{"$f":"C21*25.4","v":-91.66860000000004}]
["Y1",{"$f":"Sheet1!B78","v":-5.0597517173699664},null,{"$f":"B22*25.4","v":-128.51769362119714}]
[]
[null,null,"Downstream",null,"(PM 2)"," ","Checks"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B87","v":-5.047751717369966},{"$f":"Sheet1!A128","v":-7.0000000000014495E-3},{"$f":"B26*25.4","v":-128.21289362119714},{"$f":"C26*25.4","v":-0.17780000000003682},null,{"$f":"H13-D27","v":156.81959999999992},{"$f":"C13-E26","v":88.519000000000005}]
["X1",{"$f":"Sheet1!A87","v":-1.454751717369966},{"$f":"Sheet1!B128","v":-3.6090000000000018},{"$f":"B27*25.4","v":-36.950693621197132},{"$f":"C27*25.4","v":-91.66860000000004}]
["Y2",{"$f":"Sheet1!C87","v":-5.0517517173699673},null,{"$f":"B28*25.4","v":-128.31449362119716}]
[]
[]
[]
[]
[null,"Probe",null,null,null,"Flats"]
[]
[null,"X (in)","X (mm)",null," ","Y(in)","Y(mm)"]
[null,{"$f":"Sheet1!C141","v":-5.5529999999999999},{"$f":"B36*25.4","v":-141.0462},null,"US  +",{"$f":"Sheet1!A197","v":1.7099999999999991},{"$f":"F36*25.4","v":43.433999999999976}]
[null," ",null,null,"US   -",{"$f":"Sheet1!B197","v":1.7109999999999985},{"$f":"F37*25.4","v":43.45939999999996}]
[null,null,null,null,"DS  +",{"$f":"Sheet1!C197","v":1.7079999999999984},{"$f":"F38*25.4","v":43.38319999999996}]
[null,null,"Top of Gage Block",null,"DS   -",{"$f":"Sheet1!D197","v":1.7089999999999979},{"$f":"F39*25.4","v":43.408599999999943}]
[]
[null,"Y (in)","Y (mm)"]
[null,{"$f":"Sheet1!A139","v":-4.1330000000000009},{"$f":"B42*25.4","v":-104.97820000000002}]

{"sheet":"Sheet3","index":2,"n_rows":52,"n_cols":9,"rows_kept":52,"cut":false,"hidden":false,"merged":[]}
[]
["Undulator"," ",13]
["Date"," ",{"$f":"Sheet1!B6","v":"2015-07-02"}]
["Crew"," ","MR,BR"]
[]
[]
[null,null,"Undulator Values"]
[]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"Probe"]
[1,{"$f":"Sheet1!A28","v":-5.0587517173699688},{"$f":"Sheet1!A169","v":3.4819999999999984},{"$f":"SUM(B10*25.4)","v":-128.49229362119721},{"$f":"SUM(C10*25.4)","v":88.442799999999949}]
[2,{"$f":"Sheet1!B28","v":-5.0517517173699673},{"$f":"Sheet1!B169","v":3.4819999999999984},{"$f":"SUM(B11*25.4)","v":-128.31449362119716},{"$f":"SUM(C11*25.4)","v":88.442799999999949},null,"X (in)","X (mm)"]
[3,{"$f":"Sheet1!C28","v":-5.0557517173699686},{"$f":"Sheet1!C169","v":3.4759999999999991},-128.41609362119721,88.290399999999977,null,{"$f":"Sheet1!C141","v":-5.5529999999999999},{"$f":"SUM(G12*25.4)","v":-141.0462}]
[4,{"$f":"Sheet1!D28","v":-5.0587517173699688},{"$f":"Sheet1!D169","v":3.4779999999999989},-128.49229362119721,88.341199999999972]
[5,{"$f":"Sheet1!A35","v":4.7192482826300317},{"$f":"Sheet1!A177","v":-3.6100000000000012},119.86890637880281,-91.694000000000031]
[6,{"$f":"Sheet1!B35","v":4.7172482826300328},{"$f":"Sheet1!B177","v":-3.615000000000002},119.81810637880282,-91.821000000000041,null,"Top of Gage Block"]
[7,{"$f":"Sheet1!C35","v":4.7182482826300323},{"$f":"Sheet1!C177","v":-3.6140000000000008},119.84350637880281,-91.795600000000007]
[8,{"$f":"Sheet1!D35","v":4.7192482826300317},{"$f":"Sheet1!D177","v":-3.6190000000000015},119.86890637880281,-91.922600000000031,null,"Y (in)","Y (mm)"]
[null,null,null,null,null,null,{"$f":"Sheet1!A139","v":-4.1330000000000009},{"$f":"SUM(G18*25.4)","v":-104.97820000000002}]
[null,null,"Undulator PM Values"]
[]
[null,null,"Upstream"," ","(PM 3)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"Flats"]
["Y2",{"$f":"Sheet1!A78","v":-1.4657517173699688},{"$f":"Sheet1!C118","v":-6.0000000000020037E-3},{"$f":"SUM(B23*25.4)","v":-37.230093621197206},{"$f":"SUM(C23*25.4)","v":-0.15240000000005088}]
["X1",{"$f":"Sheet1!A78","v":-1.4657517173699688},{"$f":"Sheet1!D118","v":-3.6090000000000018},{"$f":"SUM(B24*25.4)","v":-37.230093621197206},{"$f":"SUM(C24*25.4)","v":-91.66860000000004},null," ","Y(in)","Y(mm)"]
["Y1",{"$f":"Sheet1!B78","v":-5.0597517173699664},null,{"$f":"SUM(B25*25.4)","v":-128.51769362119714},null,null,"US  +",{"$f":"Sheet1!A197","v":1.7099999999999991},{"$f":"SUM(H25*25.4)","v":43.433999999999976}]
[null,null,null,null,null,null,"US   -",{"$f":"Sheet1!B197","v":1.7109999999999985},{"$f":"SUM(H26*25.4)","v":43.45939999999996}]
[null,null,"Downstream",null,"(PM 2)",null,"DS  +",{"$f":"Sheet1!C197","v":1.7079999999999984},{"$f":"SUM(H27*25.4)","v":43.38319999999996}]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"DS   -",{"$f":"Sheet1!D197","v":1.7089999999999979},{"$f":"SUM(H28*25.4)","v":43.408599999999943}]
["Y1",{"$f":"Sheet1!B87","v":-5.047751717369966},{"$f":"Sheet1!A128","v":-7.0000000000014495E-3},{"$f":"SUM(B29*25.4)","v":-128.21289362119714},{"$f":"SUM(C29*25.4)","v":-0.17780000000003682}]
["X1",{"$f":"Sheet1!A87","v":-1.454751717369966},{"$f":"Sheet1!B128","v":-3.6090000000000018},{"$f":"SUM(B30*25.4)","v":-36.950693621197132},{"$f":"SUM(C30*25.4)","v":-91.66860000000004}]
["Y2",{"$f":"Sheet1!C87","v":-5.0517517173699673},null,{"$f":"SUM(B31*25.4)","v":-128.31449362119716}]
[]
[null,null,"Ref. PM Values"]
[]
[null,null,"Upstream"," ","(PM 4)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B68","v":-3.5997517173699691},{"$f":"Sheet1!A118","v":4.9999999999990052E-3},{"$f":"SUM(B37*25.4)","v":-91.433693621197207},{"$f":"SUM(C37*25.4)","v":0.12699999999997472},null,null," "," "]
["X1",0,{"$f":"Sheet1!B118","v":-3.5970000000000013},0,{"$f":"SUM(C38*25.4)","v":-91.363800000000026}]
["Y2",{"$f":"Sheet1!C68","v":-3.5987517173699679},null,{"$f":"SUM(B39*25.4)","v":-91.408293621197174}]
[]
[null,null,"Downstream",null,"(PM 1)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B96","v":-3.5987517173699679},{"$f":"Sheet1!C128","v":-5.0000000000007816E-3},{"$f":"SUM(B43*25.4)","v":-91.408293621197174},{"$f":"SUM(C43*25.4)","v":-0.12700000000001985}]
["X1",0,{"$f":"Sheet1!D128","v":-3.6060000000000016},0,{"$f":"SUM(C44*25.4)","v":-91.59240000000004}]
["Y2",{"$f":"Sheet1!C96","v":-3.5927517173699677},null,{"$f":"SUM(B45*25.4)","v":-91.255893621197174}]
[]
[null,null,"Ref. Block"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["A",{"$f":"Sheet1!A45","v":-6.6127517173699673},{"$f":"Sheet1!A188","v":-0.24400000000000155},{"$f":"SUM(B49*25.4)","v":-167.96389362119717},{"$f":"SUM(C49*25.4)","v":-6.1976000000000386}]
["B",{"$f":"Sheet1!B45","v":-2.0817517173699684},null,{"$f":"SUM(B50*25.4)","v":-52.876493621197199}]
["C",{"$f":"Sheet1!C45","v":-0.32275171736996811},{"$f":"Sheet1!C188","v":-2.3870000000000005},{"$f":"SUM(B51*25.4)","v":-8.1978936211971902},{"$f":"SUM(C51*25.4)","v":-60.62980000000001}]
["D",{"$f":"Sheet1!D45","v":-0.32275171736996811},{"$f":"Sheet1!D188","v":-5.4680000000000017},{"$f":"SUM(B52*25.4)","v":-8.1978936211971902},{"$f":"SUM(C52*25.4)","v":-138.88720000000004}]