{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"not_rendered":{"charts":0,"images":0,"macros":1,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":197,"n_cols":7,"rows_kept":197,"cut":false,"hidden":false,"merged":[]}
["MMF",null,null,"PM Checks"]
[]
["\"X\""]
[]
["Undulator","SN 13"]
["Date:","2016-01-08"]
["Crew:","HI, BR"]
[]
["BUCK IN "]
["Far Value",0,"PM1 X1"]
["Near Value",0,"PM 4 X1",null,null,{"$f":"D18+3.602","v":13.699}]
["1st Far",31,null,"14+1"]
["1st. Near",30.103000000000002,null,"10.509+1",null,"Far Blade",3.5129999999999999]
["2nd. Far",28,null,"8+1",null,"Near Blade"]
["2nd Near",29.244,null,"8.781+1"]
["Far Should Read",{"$f":"BuckIn(B10,B11,B12,B13,B14,B15)","v":29.74311069593648},null,"10.097-1",null,"Buck in",18.600000000000001]
["Near Should Read",29.74311069593648]
["LOS",29.74311069593648,null,10.097]
[null,null,null,null,null,null,{"$f":"G16-G13","v":15.087000000000002}]
["Undulator ",null,null,null,null,null,{"$f":"D18-G19","v":-4.990000000000002}]
[]
[]
["1)","2)","3)","4)"]
[33.847000000000001,33.843000000000004,33.848999999999997,33.853999999999999]
[1,1,1,1]
[{"$f":"SUM(A24:A25)","v":34.847000000000001},{"$f":"SUM(B24:B25)","v":34.843000000000004},{"$f":"SUM(C24:C25)","v":34.848999999999997},{"$f":"SUM(D24:D25)","v":34.853999999999999}]
[{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648}]
[{"$f":"SUM(A27-A26)","v":-5.1038893040635216},{"$f":"SUM(B27-B26)","v":-5.0998893040635238},{"$f":"SUM(C27-C26)","v":-5.1058893040635169},{"$f":"SUM(D27-D26)","v":-5.1108893040635195}]
[]
["5)","6)","7)","8)"]
[24.062000000000001,24.065000000000001,24.065999999999999,24.068000000000001]
[1,1,1,1]
[{"$f":"SUM(A31:A32)","v":25.062000000000001},{"$f":"SUM(B31:B32)","v":25.065000000000001},{"$f":"SUM(C31:C32)","v":25.065999999999999},{"$f":"SUM(D31:D32)","v":25.068000000000001}]
[{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648}]
[{"$f":"SUM(A34-A33)","v":4.6811106959364785},{"$f":"SUM(B34-B33)","v":4.6781106959364784},{"$f":"SUM(C34-C33)","v":4.6771106959364808},{"$f":"SUM(D34-D33)","v":4.6751106959364783}]
[]
[]
[null,"Reference","Block"]
[]
["A)","B)","C)","D)"]
[" "," "," "," "]
[1,1,1,1]
[{"$f":"SUM(A41:A42)","v":1},{"$f":"SUM(B41:B42)","v":1},{"$f":"SUM(C41:C42)","v":1},{"$f":"SUM(D41:D42)","v":1}]
[{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648}]
[{"$f":"SUM(A44-A43)","v":28.74311069593648},{"$f":"SUM(B44-B43)","v":28.74311069593648},{"$f":"SUM(C44-C43)","v":28.74311069593648},{"$f":"SUM(D44-D43)","v":28.74311069593648}]
[]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"X\""]
[]
[{"$f":"B18","v":29.74311069593648},"LOS"]
["Undulator",{"$f":"B5","v":"SN 13"}]
["Date:",{"$f":"B6","v":"2016-01-08"}]
["Crew:",{"$f":"B7","v":"HI, BR"}]
[]
[]
[null,"PM4"]
[]
["X1) Buck In","Y1)","Y2)"]
[null,32.344999999999999,32.345999999999997]
[null,1,1]
[null,{"$f":"SUM(B64:B65)","v":33.344999999999999},{"$f":"SUM(C64:C65)","v":33.345999999999997}]
[null,{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648}]
[0,{"$f":"SUM(B67-B66)","v":-3.6018893040635191},{"$f":"SUM(C67-C66)","v":-3.6028893040635168}]
[]
[]
[null,"PM3"]
[]
["X1)","Y1)","Y2)"]
[30.245999999999999,33.840000000000003,33.844999999999999]
[1,1,1]
[{"$f":"SUM(A74:A75)","v":31.245999999999999},{"$f":"SUM(B74:B75)","v":34.840000000000003},{"$f":"SUM(C74:C75)","v":34.844999999999999}]
[{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648}]
[{"$f":"SUM(A77-A76)","v":-1.5028893040635189},{"$f":"SUM(B77-B76)","v":-5.0968893040635237},{"$f":"SUM(C77-C76)","v":-5.1018893040635191}]
[]
[]
[null,"PM 2"]
["X1)","Y1)","Y2)"]
[30.245000000000001,33.843000000000004,33.841999999999999]
[1,1,1]
[{"$f":"SUM(A83:A84)","v":31.245000000000001},{"$f":"SUM(B83:B84)","v":34.843000000000004},{"$f":"SUM(C83:C84)","v":34.841999999999999}]
[{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648}]
[{"$f":"SUM(A86-A85)","v":-1.5018893040635213},{"$f":"SUM(B86-B85)","v":-5.0998893040635238},{"$f":"SUM(C86-C85)","v":-5.098889304063519}]
[]
[]
[null,"PM1"]
["X1) Buck In","Y1)","Y2)"]
[null,32.344000000000001,32.343000000000004]
[null,1,1]
[null,{"$f":"SUM(B92:B93)","v":33.344000000000001},{"$f":"SUM(C92:C93)","v":33.343000000000004}]
[null,{"$f":"B18","v":29.74311069593648},{"$f":"B18","v":29.74311069593648}]
[0,{"$f":"SUM(B95-B94)","v":-3.6008893040635215},{"$f":"SUM(C95-C94)","v":-3.5998893040635238}]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"Y\"",null,"PM4Y1","PM1Y1"]
[null,null,5.2039999999999997,5.2039999999999997]
[null,null,1,1]
["Undulator",{"$f":"B5","v":"SN 13"},{"$f":"SUM(C104+C105)","v":6.2039999999999997},{"$f":"SUM(D104+D105)","v":6.2039999999999997}]
["Date:",{"$f":"B6","v":"2016-01-08"}]
["Crew:",{"$f":"B7","v":"HI, BR"}]
[null,null,"HI =",{"$f":"AVERAGE(C106,D106)","v":6.2039999999999997}]
[]
["PM4",null,"PM3"]
[]
["Y1)","X1)","Y2)","X1"]
[{"$f":"C104","v":5.2039999999999997},8.8079999999999998,5.2190000000000003,8.8190000000000008]
[1,1,1,1]
[{"$f":"SUM(A114:A115)","v":6.2039999999999997},{"$f":"SUM(B114:B115)","v":9.8079999999999998},{"$f":"SUM(C114:C115)","v":6.2190000000000003},{"$f":"SUM(D114:D115)","v":9.8190000000000008}]
[{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997}]
[{"$f":"SUM(A117-A116)","v":0},{"$f":"SUM(B117-B116)","v":-3.6040000000000001},{"$f":"SUM(C117-C116)","v":-1.5000000000000568E-2},{"$f":"SUM(D117-D116)","v":-3.6150000000000011}]
[]
[]
["PM2",null,"PM1"]
[]
["Y1)","X1)","Y1)","X1)"]
[5.2110000000000003,8.8109999999999999,{"$f":"D104","v":5.2039999999999997},8.8079999999999998]
[1,1,1,1]
[{"$f":"SUM(A124:A125)","v":6.2110000000000003},{"$f":"SUM(B124:B125)","v":9.8109999999999999},{"$f":"SUM(C124:C125)","v":6.2039999999999997},{"$f":"SUM(D124:D125)","v":9.8079999999999998}]
[{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997}]
[{"$f":"SUM(A127-A126)","v":-7.0000000000005613E-3},{"$f":"SUM(B127-B126)","v":-3.6070000000000002},{"$f":"SUM(C127-C126)","v":0},{"$f":"SUM(D127-D126)","v":-3.6040000000000001}]
[]
[]
[]
[]
[]
[]
["Top Of Gage Block",null,"Probe Location in (X)"]
[]
[10.616,null,-4.99,"LOS"]
[{"$f":"D109","v":6.2039999999999997},null,-3.6999999999999998E-2,"Mic To Plus Edge"]
[{"$f":"SUM(A138-A137)","v":-4.4119999999999999},null,{"$f":"SUM(C137+C138)","v":-5.0270000000000001}]
[null,null,-3.9E-2]
[1.4430000000000001,null,{"$f":"(C139+C140)","v":-5.0659999999999998},"Probe Location"]
[2.2509999999999999]
[{"$f":"A141-A142","v":-0.80799999999999983}]
[{"$f":"A143+B136","v":-0.80799999999999983}]
[{"$f":"A138-A144","v":7.0119999999999996}]
[]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"Y\""]
[]
["Undulator",{"$f":"B5","v":"SN 13"}]
["Date:","2016-01-08"]
["Crew:",{"$f":"B7","v":"HI, BR"}]
[]
[]
["HI =",{"$f":"D109","v":6.2039999999999997}]
[]
["Undulator"]
[]
["1)","2)","3)","4)"]
[1.73,1.7290000000000001,1.73,1.7270000000000001]
[1,1,1,1]
[{"$f":"SUM(A165:A166)","v":2.73},{"$f":"SUM(B165:B166)","v":2.7290000000000001},{"$f":"SUM(C165:C166)","v":2.73},{"$f":"SUM(D165:D166)","v":2.7270000000000003}]
[{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997}]
[{"$f":"SUM(A168-A167)","v":3.4739999999999998},{"$f":"SUM(B168-B167)","v":3.4749999999999996},{"$f":"SUM(C168-C167)","v":3.4739999999999998},{"$f":"SUM(D168-D167)","v":3.4769999999999994}]
[]
[]
["5)","6)","7)","8)"]
[8.8089999999999993,8.8149999999999995,8.8119999999999994,8.8140000000000001]
[1,1,1,1]
[{"$f":"SUM(A173+A174)","v":9.8089999999999993},{"$f":"SUM(B173+B174)","v":9.8149999999999995},{"$f":"SUM(C173+C174)","v":9.8119999999999994},{"$f":"SUM(D173+D174)","v":9.8140000000000001}]
[{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997}]
[{"$f":"SUM(A176-A175)","v":-3.6049999999999995},{"$f":"SUM(B176-B175)","v":-3.6109999999999998},{"$f":"SUM(C176-C175)","v":-3.6079999999999997},{"$f":"SUM(D176-D175)","v":-3.6100000000000003}]
[]
[]
[]
["Reference Block"]
[]
["A)","B)","C)","D)"]
[" ",null," "," "]
[1,null,1,1]
[{"$f":"SUM(A184+A185)","v":{"$error":"#VALUE!"}},null,{"$f":"SUM(C184+C185)","v":{"$error":"#VALUE!"}},{"$f":"SUM(D184+D185)","v":{"$error":"#VALUE!"}}]
[{"$f":"D109","v":6.2039999999999997},null,{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997}]
[{"$f":"SUM(A187-A186)","v":{"$error":"#VALUE!"}},0,{"$f":"SUM(C187-C186)","v":{"$error":"#VALUE!"}},{"$f":"SUM(D187-D186)","v":{"$error":"#VALUE!"}}]
[]
[]
["Flats"]
[]
["U/S",null,"D/S"]
["(+)","(-)","(+)","(-)"]
[4.5019999999999998,4.5039999999999996,4.4950000000000001,4.4960000000000004]
[{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997},{"$f":"D109","v":6.2039999999999997}]
[1.704,1.7030000000000001,{"$f":"SUM(C196-C195)","v":1.7089999999999996},{"$f":"SUM(D196-D195)","v":1.7079999999999993}]

{"sheet":"Sheet2","index":1,"n_rows":42,"n_cols":8,"rows_kept":42,"cut":false,"hidden":false,"merged":[]}
[]
["Undulator"," ",{"$f":"Sheet1!B5","v":"SN 13"}]
["Date"," ",{"$f":"Sheet1!B6","v":"2016-01-08"}]
["Crew"," ",{"$f":"Sheet1!B7","v":"HI, BR"}]
[]
[]
["Undulator Values"]
[]
["TB","Y (in)","Y (mm)",null,null,"TB","X (in)","X (mm)"]
[1,{"$f":"Sheet1!A169","v":3.4739999999999998},{"$f":"B10*25.4","v":88.239599999999996},null,null,5,{"$f":"Sheet1!A35","v":4.6811106959364785},{"$f":"G10*25.4","v":118.90021167678655}]
[2,{"$f":"Sheet1!B169","v":3.4749999999999996},{"$f":"B11*25.4","v":88.264999999999986},null,null,6,{"$f":"Sheet1!B35","v":4.6781106959364784},{"$f":"G11*25.4","v":118.82401167678654}]
[3,{"$f":"Sheet1!C169","v":3.4739999999999998},{"$f":"B12*25.4","v":88.239599999999996},null,null,7,{"$f":"Sheet1!C35","v":4.6771106959364808},{"$f":"G12*25.4","v":118.79861167678661}]
[4,{"$f":"Sheet1!D169","v":3.4769999999999994},{"$f":"B13*25.4","v":88.315799999999982},null,null,8,{"$f":"Sheet1!D35","v":4.6751106959364783},{"$f":"G13*25.4","v":118.74781167678654}]
[]
[]
["PM Values"]
[]
[null,null,"Upstream"," ","(PM 3)",null,"Checks"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"X (mm)","Y (mm)"]
["Y2",{"$f":"Sheet1!C78","v":-5.1018893040635191},{"$f":"Sheet1!C118","v":-1.5000000000000568E-2},{"$f":"B20*25.4","v":-129.58798832321338},{"$f":"C20*25.4","v":-0.38100000000001444},null,{"$f":"H10-D21","v":157.07359999999994},{"$f":"C10-E20","v":88.62060000000001}]
["X1",{"$f":"Sheet1!A78","v":-1.5028893040635189},{"$f":"Sheet1!D118","v":-3.6150000000000011},{"$f":"B21*25.4","v":-38.173388323213381},{"$f":"C21*25.4","v":-91.821000000000026}]
["Y1",{"$f":"Sheet1!B78","v":-5.0968893040635237},null,{"$f":"B22*25.4","v":-129.46098832321348}]
[]
[null,null,"Downstream",null,"(PM 2)"," ","Checks"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B87","v":-5.0998893040635238},{"$f":"Sheet1!A128","v":-7.0000000000005613E-3},{"$f":"B26*25.4","v":-129.53718832321348},{"$f":"C26*25.4","v":-0.17780000000001425},null,{"$f":"H13-D27","v":156.89579999999998},{"$f":"C13-E26","v":88.493600000000001}]
["X1",{"$f":"Sheet1!A87","v":-1.5018893040635213},{"$f":"Sheet1!B128","v":-3.6070000000000002},{"$f":"B27*25.4","v":-38.14798832321344},{"$f":"C27*25.4","v":-91.617800000000003}]
["Y2",{"$f":"Sheet1!C87","v":-5.098889304063519},"\\",{"$f":"B28*25.4","v":-129.51178832321338}]
[]
[]
[]
[]
[null,"Probe",null,null,null,"Flats"]
[]
[null,"X (in)","X (mm)",null," ","Y(in)","Y(mm)"]
[null,{"$f":"Sheet1!C141","v":-5.0659999999999998},{"$f":"B36*25.4","v":-128.6764},null,"US  +",{"$f":"Sheet1!A197","v":1.704},{"$f":"F36*25.4","v":43.281599999999997}]
[null," ",null,null,"US   -",{"$f":"Sheet1!B197","v":1.7030000000000001},{"$f":"F37*25.4","v":43.2562}]
[null,null,null,null,"DS  +",{"$f":"Sheet1!C197","v":1.7089999999999996},{"$f":"F38*25.4","v":43.408599999999986}]
[null,null,"Top of Gage Block",null,"DS   -",{"$f":"Sheet1!D197","v":1.7079999999999993},{"$f":"F39*25.4","v":43.383199999999981}]
[]
[null,"Y (in)","Y (mm)"]
[null,{"$f":"Sheet1!A139","v":-4.4119999999999999},{"$f":"B42*25.4","v":-112.06479999999999}]

{"sheet":"Sheet3","index":2,"n_rows":52,"n_cols":9,"rows_kept":52,"cut":false,"hidden":false,"merged":[]}
[]
["Undulator"," ",13]
["Date"," ",{"$f":"Sheet1!B6","v":"2016-01-08"}]
["Crew"," ","MR,BR"]
[]
[]
[null,null,"Undulator Values"]
[]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"Probe"]
[1,{"$f":"Sheet1!A28","v":-5.1038893040635216},{"$f":"Sheet1!A169","v":3.4739999999999998},{"$f":"SUM(B10*25.4)","v":-129.63878832321345},{"$f":"SUM(C10*25.4)","v":88.239599999999996}]
[2,{"$f":"Sheet1!B28","v":-5.0998893040635238},{"$f":"Sheet1!B169","v":3.4749999999999996},{"$f":"SUM(B11*25.4)","v":-129.53718832321348},{"$f":"SUM(C11*25.4)","v":88.264999999999986},null,"X (in)","X (mm)"]
[3,{"$f":"Sheet1!C28","v":-5.1058893040635169},{"$f":"Sheet1!C169","v":3.4739999999999998},-129.68958832321331,88.239599999999996,null,{"$f":"Sheet1!C141","v":-5.0659999999999998},{"$f":"SUM(G12*25.4)","v":-128.6764}]
[4,{"$f":"Sheet1!D28","v":-5.1108893040635195},{"$f":"Sheet1!D169","v":3.4769999999999994},-129.81658832321338,88.315799999999982]
[5,{"$f":"Sheet1!A35","v":4.6811106959364785},{"$f":"Sheet1!A177","v":-3.6049999999999995},118.90021167678655,-91.566999999999979]
[6,{"$f":"Sheet1!B35","v":4.6781106959364784},{"$f":"Sheet1!B177","v":-3.6109999999999998},118.82401167678654,-91.719399999999993,null,"Top of Gage Block"]
[7,{"$f":"Sheet1!C35","v":4.6771106959364808},{"$f":"Sheet1!C177","v":-3.6079999999999997},118.79861167678661,-91.643199999999993]
[8,{"$f":"Sheet1!D35","v":4.6751106959364783},{"$f":"Sheet1!D177","v":-3.6100000000000003},118.74781167678654,-91.694000000000003,null,"Y (in)","Y (mm)"]
[null,null,null,null,null,null,{"$f":"Sheet1!A139","v":-4.4119999999999999},{"$f":"SUM(G18*25.4)","v":-112.06479999999999}]
[null,null,"Undulator PM Values"]
[]
[null,null,"Upstream"," ","(PM 3)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"Flats"]
["Y2",{"$f":"Sheet1!A78","v":-1.5028893040635189},{"$f":"Sheet1!C118","v":-1.5000000000000568E-2},{"$f":"SUM(B23*25.4)","v":-38.173388323213381},{"$f":"SUM(C23*25.4)","v":-0.38100000000001444}]
["X1",{"$f":"Sheet1!A78","v":-1.5028893040635189},{"$f":"Sheet1!D118","v":-3.6150000000000011},{"$f":"SUM(B24*25.4)","v":-38.173388323213381},{"$f":"SUM(C24*25.4)","v":-91.821000000000026},null," ","Y(in)","Y(mm)"]
["Y1",{"$f":"Sheet1!B78","v":-5.0968893040635237},null,{"$f":"SUM(B25*25.4)","v":-129.46098832321348},null,null,"US  +",{"$f":"Sheet1!A197","v":1.704},{"$f":"SUM(H25*25.4)","v":43.281599999999997}]
[null,null,null,null,null,null,"US   -",{"$f":"Sheet1!B197","v":1.7030000000000001},{"$f":"SUM(H26*25.4)","v":43.2562}]
[null,null,"Downstream",null,"(PM 2)",null,"DS  +",{"$f":"Sheet1!C197","v":1.7089999999999996},{"$f":"SUM(H27*25.4)","v":43.408599999999986}]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"DS   -",{"$f":"Sheet1!D197","v":1.7079999999999993},{"$f":"SUM(H28*25.4)","v":43.383199999999981}]
["Y1",{"$f":"Sheet1!B87","v":-5.0998893040635238},{"$f":"Sheet1!A128","v":-7.0000000000005613E-3},{"$f":"SUM(B29*25.4)","v":-129.53718832321348},{"$f":"SUM(C29*25.4)","v":-0.17780000000001425}]
["X1",{"$f":"Sheet1!A87","v":-1.5018893040635213},{"$f":"Sheet1!B128","v":-3.6070000000000002},{"$f":"SUM(B30*25.4)","v":-38.14798832321344},{"$f":"SUM(C30*25.4)","v":-91.617800000000003}]
["Y2",{"$f":"Sheet1!C87","v":-5.098889304063519},null,{"$f":"SUM(B31*25.4)","v":-129.51178832321338}]
[]
[null,null,"Ref. PM Values"]
[]
[null,null,"Upstream"," ","(PM 4)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B68","v":-3.6018893040635191},{"$f":"Sheet1!A118","v":0},{"$f":"SUM(B37*25.4)","v":-91.487988323213386},{"$f":"SUM(C37*25.4)","v":0},null,null," "," "]
["X1",0,{"$f":"Sheet1!B118","v":-3.6040000000000001},0,{"$f":"SUM(C38*25.4)","v":-91.541600000000003}]
["Y2",{"$f":"Sheet1!C68","v":-3.6028893040635168},null,{"$f":"SUM(B39*25.4)","v":-91.51338832321332}]
[]
[null,null,"Downstream",null,"(PM 1)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B96","v":-3.6008893040635215},{"$f":"Sheet1!C128","v":0},{"$f":"SUM(B43*25.4)","v":-91.462588323213438},{"$f":"SUM(C43*25.4)","v":0}]
["X1",0,{"$f":"Sheet1!D128","v":-3.6040000000000001},0,{"$f":"SUM(C44*25.4)","v":-91.541600000000003}]
["Y2",{"$f":"Sheet1!C96","v":-3.5998893040635238},null,{"$f":"SUM(B45*25.4)","v":-91.437188323213505}]
[]
[null,null,"Ref. Block"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["A",{"$f":"Sheet1!A45","v":28.74311069593648},{"$f":"Sheet1!A188","v":{"$error":"#VALUE!"}},{"$f":"SUM(B49*25.4)","v":730.07501167678652},{"$f":"SUM(C49*25.4)","v":{"$error":"#VALUE!"}}]
["B",{"$f":"Sheet1!B45","v":28.74311069593648},null,{"$f":"SUM(B50*25.4)","v":730.07501167678652}]
["C",{"$f":"Sheet1!C45","v":28.74311069593648},{"$f":"Sheet1!C188","v":{"$error":"#VALUE!"}},{"$f":"SUM(B51*25.4)","v":730.07501167678652},{"$f":"SUM(C51*25.4)","v":{"$error":"#VALUE!"}}]
["D",{"$f":"Sheet1!D45","v":28.74311069593648},{"$f":"Sheet1!D188","v":{"$error":"#VALUE!"}},{"$f":"SUM(B52*25.4)","v":730.07501167678652},{"$f":"SUM(C52*25.4)","v":{"$error":"#VALUE!"}}]