{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"not_rendered":{"charts":0,"images":0,"macros":1,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":197,"n_cols":4,"rows_kept":197,"cut":false,"hidden":false,"merged":[]}
["MMF",null,null,"PM Checks"]
[]
["\"X\""]
[]
["Undulator","SN 31"]
["Date:","2012-05-03"]
["Crew:","BR, LG"]
[]
["BUCK IN "]
["Far Value",0,"PM1 X1"]
["Near Value",0,"PM 4 X1"]
["1st Far",27]
["1st. Near",27.116]
["2nd. Far",27.5]
["2nd Near",27.247]
["Far Should Read",{"$f":"BuckIn(B10,B11,B12,B13,B14,B15)","v":27.157181571815716}]
["Near Should Read",27.157181571815716]
["LOS",27.157181571815716]
[]
["Undulator "]
[]
[]
["1)","2)","3)","4)"]
[31.643000000000001,31.643999999999998,31.643999999999998,31.645]
[1,1,1,1]
[{"$f":"SUM(A24:A25)","v":32.643000000000001},{"$f":"SUM(B24:B25)","v":32.643999999999998},{"$f":"SUM(C24:C25)","v":32.643999999999998},{"$f":"SUM(D24:D25)","v":32.644999999999996}]
[{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716}]
[{"$f":"SUM(A27-A26)","v":-5.4858184281842846},{"$f":"SUM(B27-B26)","v":-5.4868184281842822},{"$f":"SUM(C27-C26)","v":-5.4868184281842822},{"$f":"SUM(D27-D26)","v":-5.4878184281842799}]
[]
["5)","6)","7)","8)"]
[21.861000000000001,21.864999999999998,21.866,21.866]
[1,1,1,1]
[{"$f":"SUM(A31:A32)","v":22.861000000000001},{"$f":"SUM(B31:B32)","v":22.864999999999998},{"$f":"SUM(C31:C32)","v":22.866},{"$f":"SUM(D31:D32)","v":22.866}]
[{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716}]
[{"$f":"SUM(A34-A33)","v":4.2961815718157155},{"$f":"SUM(B34-B33)","v":4.2921815718157177},{"$f":"SUM(C34-C33)","v":4.2911815718157165},{"$f":"SUM(D34-D33)","v":4.2911815718157165}]
[]
[]
[null,"Reference","Block"]
[]
["A)","B)","C)","D)"]
[33.192,28.658000000000001,26.898,26.899000000000001]
[1,1,1,1]
[{"$f":"SUM(A41:A42)","v":34.192},{"$f":"SUM(B41:B42)","v":29.658000000000001},{"$f":"SUM(C41:C42)","v":27.898},{"$f":"SUM(D41:D42)","v":27.899000000000001}]
[{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716}]
[{"$f":"SUM(A44-A43)","v":-7.034818428184284},{"$f":"SUM(B44-B43)","v":-2.5008184281842851},{"$f":"SUM(C44-C43)","v":-0.74081842818428356},{"$f":"SUM(D44-D43)","v":-0.74181842818428478}]
[]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"X\""]
[]
[{"$f":"B18","v":27.157181571815716},"LOS"]
["Undulator",{"$f":"B5","v":"SN 31"}]
["Date:",{"$f":"B6","v":"2012-05-03"}]
["Crew:",{"$f":"B7","v":"BR, LG"}]
[]
[]
[null,"PM4"]
[]
["X1) Buck In","Y1)","Y2)"]
[null,29.757999999999999,29.76]
[null,1,1]
[null,{"$f":"SUM(B64:B65)","v":30.757999999999999},{"$f":"SUM(C64:C65)","v":30.76}]
[null,{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716}]
[0,{"$f":"SUM(B67-B66)","v":-3.600818428184283},{"$f":"SUM(C67-C66)","v":-3.6028184281842854}]
[]
[]
[null,"PM3"]
[]
["X1)","Y1)","Y2)"]
[28.036000000000001,31.631,31.631]
[1,1,1]
[{"$f":"SUM(A74:A75)","v":29.036000000000001},{"$f":"SUM(B74:B75)","v":32.631},{"$f":"SUM(C74:C75)","v":32.631}]
[{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716}]
[{"$f":"SUM(A77-A76)","v":-1.8788184281842852},{"$f":"SUM(B77-B76)","v":-5.4738184281842841},{"$f":"SUM(C77-C76)","v":-5.4738184281842841}]
[]
[]
[null,"PM 2"]
["X1)","Y1)","Y2)"]
[28.047000000000001,31.640999999999998,31.641999999999999]
[1,1,1]
[{"$f":"SUM(A83:A84)","v":29.047000000000001},{"$f":"SUM(B83:B84)","v":32.640999999999998},{"$f":"SUM(C83:C84)","v":32.641999999999996}]
[{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716}]
[{"$f":"SUM(A86-A85)","v":-1.8898184281842845},{"$f":"SUM(B86-B85)","v":-5.4838184281842821},{"$f":"SUM(C86-C85)","v":-5.4848184281842798}]
[]
[]
[null,"PM1"]
["X1) Buck In","Y1)","Y2)"]
[null,29.759,29.754000000000001]
[null,1,1]
[null,{"$f":"SUM(B92:B93)","v":30.759},{"$f":"SUM(C92:C93)","v":30.754000000000001}]
[null,{"$f":"B18","v":27.157181571815716},{"$f":"B18","v":27.157181571815716}]
[0,{"$f":"SUM(B95-B94)","v":-3.6018184281842842},{"$f":"SUM(C95-C94)","v":-3.5968184281842852}]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"Y\"",null,"PM4Y1","PM1Y1"]
[null,null,8.3160000000000007,8.3059999999999992]
[null,null,1,1]
["Undulator",{"$f":"B5","v":"SN 31"},{"$f":"SUM(C104+C105)","v":9.3160000000000007},{"$f":"SUM(D104+D105)","v":9.3059999999999992}]
["Date:",{"$f":"B6","v":"2012-05-03"}]
["Crew:",{"$f":"B7","v":"BR, LG"}]
[null,null,"HI =",{"$f":"AVERAGE(C106,D106)","v":9.3109999999999999}]
[]
["PM4",null,"PM3"]
[]
["Y1)","X1)","Y2)","X1"]
[{"$f":"C104","v":8.3160000000000007},11.925000000000001,8.32,11.920999999999999]
[1,1,1,1]
[{"$f":"SUM(A114:A115)","v":9.3160000000000007},{"$f":"SUM(B114:B115)","v":12.925000000000001},{"$f":"SUM(C114:C115)","v":9.32},{"$f":"SUM(D114:D115)","v":12.920999999999999}]
[{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999}]
[{"$f":"SUM(A117-A116)","v":-5.0000000000007816E-3},{"$f":"SUM(B117-B116)","v":-3.6140000000000008},{"$f":"SUM(C117-C116)","v":-9.0000000000003411E-3},{"$f":"SUM(D117-D116)","v":-3.6099999999999994}]
[]
[]
["PM2",null,"PM1"]
[]
["Y1)","X1)","Y1)","X1)"]
[8.3190000000000008,11.92,{"$f":"D104","v":8.3059999999999992},11.907999999999999]
[1,1,1,1]
[{"$f":"SUM(A124:A125)","v":9.3190000000000008},{"$f":"SUM(B124:B125)","v":12.92},{"$f":"SUM(C124:C125)","v":9.3059999999999992},{"$f":"SUM(D124:D125)","v":12.907999999999999}]
[{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999}]
[{"$f":"SUM(A127-A126)","v":-8.0000000000008953E-3},{"$f":"SUM(B127-B126)","v":-3.609},{"$f":"SUM(C127-C126)","v":5.0000000000007816E-3},{"$f":"SUM(D127-D126)","v":-3.5969999999999995}]
[]
[]
[]
[]
[]
[]
["Top Of Gage Block",null,"Probe Location in (X)"]
[]
[13.718,null,-5.4640000000000004,"LOS"]
[{"$f":"D109","v":9.3109999999999999},null,4.1000000000000002E-2,"Mic To Plus Edge"]
[{"$f":"SUM(A138-A137)","v":-4.407},null,{"$f":"SUM(C137+C138)","v":-5.423}]
[null,null,-3.9E-2]
[null,null,{"$f":"(C139+C140)","v":-5.4619999999999997},"Probe Location"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"Y\""]
[]
["Undulator",{"$f":"B5","v":"SN 31"}]
["Date:",{"$f":"B6","v":"2012-05-03"}]
["Crew:",{"$f":"B7","v":"BR, LG"}]
[]
[]
["HI =",{"$f":"D109","v":9.3109999999999999}]
[]
["Undulator"]
[]
["1)","2)","3)","4)"]
[4.8380000000000001,4.84,4.84,4.8360000000000003]
[1,1,1,1]
[{"$f":"SUM(A165:A166)","v":5.8380000000000001},{"$f":"SUM(B165:B166)","v":5.84},{"$f":"SUM(C165:C166)","v":5.84},{"$f":"SUM(D165:D166)","v":5.8360000000000003}]
[{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999}]
[{"$f":"SUM(A168-A167)","v":3.4729999999999999},{"$f":"SUM(B168-B167)","v":3.4710000000000001},{"$f":"SUM(C168-C167)","v":3.4710000000000001},{"$f":"SUM(D168-D167)","v":3.4749999999999996}]
[]
[]
["5)","6)","7)","8)"]
[11.916,11.914999999999999,11.917999999999999,11.917]
[1,1,1,1]
[{"$f":"SUM(A173+A174)","v":12.916},{"$f":"SUM(B173+B174)","v":12.914999999999999},{"$f":"SUM(C173+C174)","v":12.917999999999999},{"$f":"SUM(D173+D174)","v":12.917}]
[{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999}]
[{"$f":"SUM(A176-A175)","v":-3.6050000000000004},{"$f":"SUM(B176-B175)","v":-3.6039999999999992},{"$f":"SUM(C176-C175)","v":-3.6069999999999993},{"$f":"SUM(D176-D175)","v":-3.6059999999999999}]
[]
[]
[]
["Reference Block"]
[]
["A)","B)","C)","D)"]
[8.5570000000000004,null,10.704000000000001,13.784000000000001]
[1,null,1,1]
[{"$f":"SUM(A184+A185)","v":9.5570000000000004},null,{"$f":"SUM(C184+C185)","v":11.704000000000001},{"$f":"SUM(D184+D185)","v":14.784000000000001}]
[{"$f":"D109","v":9.3109999999999999},null,{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999}]
[{"$f":"SUM(A187-A186)","v":-0.24600000000000044},0,{"$f":"SUM(C187-C186)","v":-2.3930000000000007},{"$f":"SUM(D187-D186)","v":-5.4730000000000008}]
[]
[]
["Flats"]
[]
["U/S",null,"D/S"]
["(+)","(-)","(+)","(-)"]
[7.601,7.601,7.6,7.601]
[{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999},{"$f":"D109","v":9.3109999999999999}]
[{"$f":"SUM(A196-A195)","v":1.71},{"$f":"SUM(B196-B195)","v":1.71},{"$f":"SUM(C196-C195)","v":1.7110000000000003},{"$f":"SUM(D196-D195)","v":1.71}]

{"sheet":"Sheet2","index":1,"n_rows":42,"n_cols":8,"rows_kept":42,"cut":false,"hidden":false,"merged":[]}
[]
["Undulator"," ",{"$f":"Sheet1!B5","v":"SN 31"}]
["Date"," ",{"$f":"Sheet1!B6","v":"2012-05-03"}]
["Crew"," ",{"$f":"Sheet1!B7","v":"BR, LG"}]
[]
[]
["Undulator Values"]
[]
["TB","Y (in)","Y (mm)",null,null,"TB","X (in)","X (mm)"]
[1,{"$f":"Sheet1!A169","v":3.4729999999999999},{"$f":"B10*25.4","v":88.214199999999991},null,null,5,{"$f":"Sheet1!A35","v":4.2961815718157155},{"$f":"G10*25.4","v":109.12301192411917}]
[2,{"$f":"Sheet1!B169","v":3.4710000000000001},{"$f":"B11*25.4","v":88.163399999999996},null,null,6,{"$f":"Sheet1!B35","v":4.2921815718157177},{"$f":"G11*25.4","v":109.02141192411922}]
[3,{"$f":"Sheet1!C169","v":3.4710000000000001},{"$f":"B12*25.4","v":88.163399999999996},null,null,7,{"$f":"Sheet1!C35","v":4.2911815718157165},{"$f":"G12*25.4","v":108.99601192411919}]
[4,{"$f":"Sheet1!D169","v":3.4749999999999996},{"$f":"B13*25.4","v":88.264999999999986},null,null,8,{"$f":"Sheet1!D35","v":4.2911815718157165},{"$f":"G13*25.4","v":108.99601192411919}]
[]
[]
["PM Values"]
[]
[null,null,"Upstream"," ","(PM 3)",null,"Checks"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"X (mm)","Y (mm)"]
["Y2",{"$f":"Sheet1!C78","v":-5.4738184281842841},{"$f":"Sheet1!C118","v":-9.0000000000003411E-3},{"$f":"B20*25.4","v":-139.03498807588082},{"$f":"C20*25.4","v":-0.22860000000000866},null,{"$f":"H10-D21","v":156.84500000000003},{"$f":"C10-E20","v":88.442800000000005}]
["X1",{"$f":"Sheet1!A78","v":-1.8788184281842852},{"$f":"Sheet1!D118","v":-3.6099999999999994},{"$f":"B21*25.4","v":-47.721988075880844},{"$f":"C21*25.4","v":-91.693999999999974}]
["Y1",{"$f":"Sheet1!B78","v":-5.4738184281842841},null,{"$f":"B22*25.4","v":-139.03498807588082}]
[]
[null,null,"Downstream",null,"(PM 2)"," ","Checks"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B87","v":-5.4838184281842821},{"$f":"Sheet1!A128","v":-8.0000000000008953E-3},{"$f":"B26*25.4","v":-139.28898807588075},{"$f":"C26*25.4","v":-0.20320000000002272},null,{"$f":"H13-D27","v":156.99740000000003},{"$f":"C13-E26","v":88.46820000000001}]
["X1",{"$f":"Sheet1!A87","v":-1.8898184281842845},{"$f":"Sheet1!B128","v":-3.609},{"$f":"B27*25.4","v":-48.001388075880826},{"$f":"C27*25.4","v":-91.668599999999998}]
["Y2",{"$f":"Sheet1!C87","v":-5.4848184281842798},"\\",{"$f":"B28*25.4","v":-139.31438807588069}]
[]
[]
[]
[]
[null,"Probe",null,null,null,"Flats"]
[]
[null,"X (in)","X (mm)",null," ","Y(in)","Y(mm)"]
[null,{"$f":"Sheet1!C141","v":-5.4619999999999997},{"$f":"B36*25.4","v":-138.73479999999998},null,"US  +",{"$f":"Sheet1!A197","v":1.71},{"$f":"F36*25.4","v":43.433999999999997}]
[null," ",null,null,"US   -",{"$f":"Sheet1!B197","v":1.71},{"$f":"F37*25.4","v":43.433999999999997}]
[null,null,null,null,"DS  +",{"$f":"Sheet1!C197","v":1.7110000000000003},{"$f":"F38*25.4","v":43.459400000000002}]
[null,null,"Top of Gage Block",null,"DS   -",{"$f":"Sheet1!D197","v":1.71},{"$f":"F39*25.4","v":43.433999999999997}]
[]
[null,"Y (in)","Y (mm)"]
[null,{"$f":"Sheet1!A139","v":-4.407},{"$f":"B42*25.4","v":-111.9378}]

{"sheet":"Sheet3","index":2,"n_rows":52,"n_cols":9,"rows_kept":52,"cut":false,"hidden":false,"merged":[]}
[]
["Undulator"," ",13]
["Date"," ",{"$f":"Sheet1!B6","v":"2012-05-03"}]
["Crew"," ","MR,BR"]
[]
[]
[null,null,"Undulator Values"]
[]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"Probe"]
[1,{"$f":"Sheet1!A28","v":-5.4858184281842846},{"$f":"Sheet1!A169","v":3.4729999999999999},{"$f":"SUM(B10*25.4)","v":-139.33978807588082},{"$f":"SUM(C10*25.4)","v":88.214199999999991}]
[2,{"$f":"Sheet1!B28","v":-5.4868184281842822},{"$f":"Sheet1!B169","v":3.4710000000000001},{"$f":"SUM(B11*25.4)","v":-139.36518807588075},{"$f":"SUM(C11*25.4)","v":88.163399999999996},null,"X (in)","X (mm)"]
[3,{"$f":"Sheet1!C28","v":-5.4868184281842822},{"$f":"Sheet1!C169","v":3.4710000000000001},-139.36518807588075,88.163399999999996,null,{"$f":"Sheet1!C141","v":-5.4619999999999997},{"$f":"SUM(G12*25.4)","v":-138.73479999999998}]
[4,{"$f":"Sheet1!D28","v":-5.4878184281842799},{"$f":"Sheet1!D169","v":3.4749999999999996},-139.39058807588071,88.264999999999986]
[5,{"$f":"Sheet1!A35","v":4.2961815718157155},{"$f":"Sheet1!A177","v":-3.6050000000000004},109.12301192411917,-91.567000000000007]
[6,{"$f":"Sheet1!B35","v":4.2921815718157177},{"$f":"Sheet1!B177","v":-3.6039999999999992},109.02141192411922,-91.541599999999974,null,"Top of Gage Block"]
[7,{"$f":"Sheet1!C35","v":4.2911815718157165},{"$f":"Sheet1!C177","v":-3.6069999999999993},108.99601192411919,-91.617799999999974]
[8,{"$f":"Sheet1!D35","v":4.2911815718157165},{"$f":"Sheet1!D177","v":-3.6059999999999999},108.99601192411919,-91.592399999999998,null,"Y (in)","Y (mm)"]
[null,null,null,null,null,null,{"$f":"Sheet1!A139","v":-4.407},{"$f":"SUM(G18*25.4)","v":-111.9378}]
[null,null,"Undulator PM Values"]
[]
[null,null,"Upstream"," ","(PM 3)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"Flats"]
["Y2",{"$f":"Sheet1!A78","v":-1.8788184281842852},{"$f":"Sheet1!C118","v":-9.0000000000003411E-3},{"$f":"SUM(B23*25.4)","v":-47.721988075880844},{"$f":"SUM(C23*25.4)","v":-0.22860000000000866}]
["X1",{"$f":"Sheet1!A78","v":-1.8788184281842852},{"$f":"Sheet1!D118","v":-3.6099999999999994},{"$f":"SUM(B24*25.4)","v":-47.721988075880844},{"$f":"SUM(C24*25.4)","v":-91.693999999999974},null," ","Y(in)","Y(mm)"]
["Y1",{"$f":"Sheet1!B78","v":-5.4738184281842841},null,{"$f":"SUM(B25*25.4)","v":-139.03498807588082},null,null,"US  +",{"$f":"Sheet1!A197","v":1.71},{"$f":"SUM(H25*25.4)","v":43.433999999999997}]
[null,null,null,null,null,null,"US   -",{"$f":"Sheet1!B197","v":1.71},{"$f":"SUM(H26*25.4)","v":43.433999999999997}]
[null,null,"Downstream",null,"(PM 2)",null,"DS  +",{"$f":"Sheet1!C197","v":1.7110000000000003},{"$f":"SUM(H27*25.4)","v":43.459400000000002}]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"DS   -",{"$f":"Sheet1!D197","v":1.71},{"$f":"SUM(H28*25.4)","v":43.433999999999997}]
["Y1",{"$f":"Sheet1!B87","v":-5.4838184281842821},{"$f":"Sheet1!A128","v":-8.0000000000008953E-3},{"$f":"SUM(B29*25.4)","v":-139.28898807588075},{"$f":"SUM(C29*25.4)","v":-0.20320000000002272}]
["X1",{"$f":"Sheet1!A87","v":-1.8898184281842845},{"$f":"Sheet1!B128","v":-3.609},{"$f":"SUM(B30*25.4)","v":-48.001388075880826},{"$f":"SUM(C30*25.4)","v":-91.668599999999998}]
["Y2",{"$f":"Sheet1!C87","v":-5.4848184281842798},null,{"$f":"SUM(B31*25.4)","v":-139.31438807588069}]
[]
[null,null,"Ref. PM Values"]
[]
[null,null,"Upstream"," ","(PM 4)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B68","v":-3.600818428184283},{"$f":"Sheet1!A118","v":-5.0000000000007816E-3},{"$f":"SUM(B37*25.4)","v":-91.460788075880785},{"$f":"SUM(C37*25.4)","v":-0.12700000000001985},null,null," "," "]
["X1",0,{"$f":"Sheet1!B118","v":-3.6140000000000008},0,{"$f":"SUM(C38*25.4)","v":-91.795600000000007}]
["Y2",{"$f":"Sheet1!C68","v":-3.6028184281842854},null,{"$f":"SUM(B39*25.4)","v":-91.511588075880852}]
[]
[null,null,"Downstream",null,"(PM 1)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B96","v":-3.6018184281842842},{"$f":"Sheet1!C128","v":5.0000000000007816E-3},{"$f":"SUM(B43*25.4)","v":-91.486188075880818},{"$f":"SUM(C43*25.4)","v":0.12700000000001985}]
["X1",0,{"$f":"Sheet1!D128","v":-3.5969999999999995},0,{"$f":"SUM(C44*25.4)","v":-91.363799999999983}]
["Y2",{"$f":"Sheet1!C96","v":-3.5968184281842852},null,{"$f":"SUM(B45*25.4)","v":-91.359188075880837}]
[]
[null,null,"Ref. Block"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["A",{"$f":"Sheet1!A45","v":-7.034818428184284},{"$f":"Sheet1!A188","v":-0.24600000000000044},{"$f":"SUM(B49*25.4)","v":-178.6843880758808},{"$f":"SUM(C49*25.4)","v":-6.2484000000000108}]
["B",{"$f":"Sheet1!B45","v":-2.5008184281842851},null,{"$f":"SUM(B50*25.4)","v":-63.520788075880837}]
["C",{"$f":"Sheet1!C45","v":-0.74081842818428356},{"$f":"Sheet1!C188","v":-2.3930000000000007},{"$f":"SUM(B51*25.4)","v":-18.816788075880801},{"$f":"SUM(C51*25.4)","v":-60.782200000000017}]
["D",{"$f":"Sheet1!D45","v":-0.74181842818428478},{"$f":"Sheet1!D188","v":-5.4730000000000008},{"$f":"SUM(B52*25.4)","v":-18.842188075880834},{"$f":"SUM(C52*25.4)","v":-139.01420000000002}]