Magnet Measurement Test Plan

LTB/BTS quadrupole measurements – Rev 0

Logistic Info

| Person Responsible | Domenico Dell’Orco – Robert DiMattia (x5457) |
| --- | --- |
| Phone # | 650-926-8683 (cell. 510-219-0378) |
| Start Date | 07-13-2010 |
| Finish Date |  |
| Acct # for alignment | TBD |

Magnet Designation and Type

| Project Name |  |  | LTB/BTS quadrupole measurements |  |
| --- | --- | --- | --- | --- |
| Magnet Type |  |  | Focusing Quadrupole |  |
| Magnet Name/Number |  |  | Spare magnet |  |
| Serial # / PC # |  |  | N/A |  |
| Trim Coils | Yes | Trim Types |  | Solid/ indirectly cooled |
| Magnet gap or diameter (mm) |  |  |  | 60 mm |

Alignment

| Precision (0.05 mm) |  |
| --- | --- |
| Has the magnet been fiducialized | NO |
| If it has not been fiducialized, what is origin point of the magnet, what is its coordinate system and what surfaces should be used for alignment? |  |
| + Z direction of magnet (Please provide drawing) | From Trim coil leads to Main coil leads |

Special Alignment Tolerances

| X |  |
| --- | --- |
| Y |  |
| Pitch |  |
| Yaw |  |
| Roll |  |

Water and Power.

| Water Pressure | No cooling required (keep main current ≤10 A, trim ≤1.5 A). Monitor coil temp, shut magnet if temp >45C | PSI |
| --- | --- | --- |
| Water Flow | NONE | Gpm |
| Maximum Current Main | 10 | Amps |
| Max Current Trims | 1.5 | Amps |

Polarity

| Main Coil Polarity. (Nearest North pole to Horizontal look in +z direction) | - Focusing quadrupole |
| --- | --- |
| 1st Trim Polarity (Nearest North pole to Horizontal look in +z direction) | Vertical field dipole and Horizontal field dipole |
| 2nd Trim Polarity (Nearest North pole to Horizontal look in +z direction) | N/A |

Temperature Tests

| Temperature and Resistance Test | No, Protect magnet using thermostats |
| --- | --- |
| Areas to be measured |  |
| Currents (Amps) |  |
| Check water flow |  |

Standardization

| Magnet warm up. ? time (mins) @ ? current (Amps) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Conditioning Standardization Cycles Performed before measurements: |  |  |  |  |
| Number of Stdz cycles |  |  |  |  |
| Ramp Rate (A/s) | 10 A/s | Minimum Current (A) |  | 0 |
| Settle Time (sec) | 30 s | Maximum Current (A) |  | 10A |
| Ramp Type Linear or Three Linear or Cosine |  |  | Linear |  |

Measurements

| Integrated Strength | Main | Yes |  | Trims | Yes |  |
| --- | --- | --- | --- | --- | --- | --- |
| Currents: |  |  |  |  |  |  |
| Harmonics | Main |  | Yes | Trims |  | Yes |
| Currents: |  |  |  |  |  |  |

Harmonic Tolerances (as % of main, Main =100 %)

| Dipole n = 1 | Quadrupole n = 2 | Sextupole n = 3 | Octupole n =4 |
| --- | --- | --- | --- |

Field Map

| Type of map? (Transverse, axial, 3D, Locate Center) |  |
| --- | --- |
| Dimensions ( X, Y, Z) |  |
| Spacing between step (X, Y, Z) |  |

Notes and Additional Tests:

1. Measure the quadrupole magnetic length at 10A.

2. Hysteresis Loop: measure integrated gradient with I from 0A to +10A, then to -10A, then to 10A in steps of 1 A

3.1. Repeat standardization and measurements 3 times

Standardization #1: Start from 0A , raise to 10 A, decrease to -10A, then go back to 0 A. Repeat standardization 3 times.

Measure  integrated gradient at I=0A, 2A, 4A, 6A, 8A, 10A.

3.2. Repeat standardization and measurements 3 times

Standardization #2: Start from 0A , raise to 10 A, decrease to -10A, then go back to 0 A. Repeat standardization 5 times.

Measure  integrated gradient at I=0A, 2A, 4A, 6A, 8A, 10A.

3.3. Repeat standardization and measurements 3 times

Standardization #3: Start from 0A , decrease to -10 A, raise to 10A, then go back to 0 A. Repeat standardization 3 times.

Measure  integrated gradient at I=0A, -2A,- 4A,- 6A,- 8A, -10A.

3.4. Repeat standardization and measurements 3 times

Standardization #4: Start from 0A , decrease to -10 A, raise to 10A, then go back to 0 A. Repeat standardization 5 times.

Measure  integrated gradient at I=0A, -2A,- 4A, -6A,- 8A, -10A.

3.5. Repeat standardization and measurements 3 times

Standardization #5: Start from 0A , raise to 2 A then go back to 0 A. Repeat standardization 3 times.

Measure  integrated gradient at I=0A, 2A, 4A, 6A, 8A, 10A.

3.6. Repeat standardization and measurements 3 times

Standardization #6: Start from 0A , raise to 2 A then go back to 0 A. Repeat standardization 5 times.

Measure  integrated gradient at I=0A, 2A, 4A, 6A, 8A, 10A.

3.7. Repeat standardization and measurements 3 times

Standardization #7: Start from 0A , raise to 6 A then go back to 0 A. Repeat standardization 3 times.

Measure  integrated gradient at I=0A, 2A, 4A, 6A, 8A, 10A.

3.8. Repeat standardization and measurements 3 times

Standardization #8: Start from 0A , raise to 6 A then go back to 0 A. Repeat standardization 5 times.

Measure  integrated gradient at I=0A, 2A, 4A, 6A, 8A, 10A.

4.1 Measure ∫BdL secondary coils wired as vertical dipole at I=-1.5A, -1A, -0.5A, 0.5A, 1A, 1.5A,

4.2 Measure ∫BdL secondary coils wired as horizontal dipole at I=-1.5A, -1A, -0.5A, 0.5A, 1A, 1.5A