NLCTA Safety Audit

Table of Contents:

3Background:


3Contemporary programs:


3Beam Programs


3Rf processing programs


3Future programs


4Project Organization:


4Safety Organization:


5ESH Participation in the program


5Proposed facility developments:


5Near term accelerator beamline development:


5Near term accelerator support environment development:


5Components of the NLCTA


5The accelerator - beamline


6The accelerator – support equipment


6Operation of the NLCTA


6Procedures and Operations Documents


7Operations Directives


7Beam Authorization Sheet (BAS)


7Daily Inspection Checklist


7Weekly Inspection Checklist


7Training and Operator Control


8NLCTA Safety Systems


8Radiation Safety Systems


8PPS


9BCS Safety System


9Fire Safety System


10General Safety Concerns and Areas of Interest


10Environmental Safety & Hazardous Materials


10Pressure and Vacuum Vessel


11Laser Safety


11Seismic Safety


12Crane Safety:


12Radiation Safety


12Electrical Safety, AC


12Grounding


13Multiple power sources


13Bakeout


13Electrical Safety, DC


13Occupational Safety


13NLCTA Shop


14NLCTA Storage Buildings


14ADA Accessibility





Background: 

The NLCTA was constructed (when in 1994-1995) to demonstrate beam loading compensation schemes for high current long (> fill time) beams with high (>25%) beam loading, and as a development and demonstration x-band linear accelerator.  These goals have been largely achieved.

Contemporary programs:

Recent use of the facility by NLC has been to study the lifetime and performance of X-band accelerator structures running at high gradient.  This program requires subjecting the structures to several hundred hours of operation, with brief periods of beam operation to diagnose structure operation.  This operation is allowed under the conditions approved by the Radiation Safety Officer as “unattended operation without beam,” and proceeds though the night without a fully manned control room.

Beam Programs

Using the beam, two families of structure tests are interesting.  The most common (1/2 day every week) consists of accelerating a short (1.5 nS) beam through the structure.  The beam loading of the structure leaves energy in the cavity which can be directly measured and transformed into an equivalent in-situ “bead-pull” of the cavity installed on beamline.  This is done without applying rf power to the structures under test.  

Less frequently, the direct energy gain of the structures is measured by accelerating a beam and measuring the energy gain.  The relative phasing error of structure pairs is measured using the “bead-pull” technique and the energy gain result are corrected based on this measurement.

Rf processing programs

Recent results from NLCTA indicate that accelerator structures can be produced that, following initial processing, meet or exceed the program goals for the NLC.  Unfortunately, continued operation at the NLCTA has shown us that continued operation of the structures can result in continuing damage that detunes the accelerator and destroys its effectiveness as an accelerator.

The current program emphasizes the continued operation of a number of differing structure designs to determine the best design, operation, and manufacturing methods to produce long-life structures.  Continued rf processing is occasionally interrupted to allow an in-situ bead pull to monitor performance.  Run lengths of 500 to 2000 processing hours are desired.

Future programs

There are several facility development programs proceeding simultaneously.

The NLC project needs to develop and demonstrate reliable production of high power rf (goal of around 600 MW by 400 nS delivered to 8 loads sequentially.  This 8-pack
 project will proceed in several phases: (1) the installation of a 4-klystron modulator and high power SLED-II rf folded delay line compression, followed by the namesake project of an 8-klystron modulator (the “8-pack”) followed by an overmoded delay-line distribution system which allows the 8 klystron’s power to be summed to 8 output ports sequentially.  One or two output ports will eventually be powering accelerator structures in the NLCTA.

The ORION
 project and the implications to the Operation and Safety of the NLCTA will not be discussed.  This project is not an approved experimental program at present, and will require a thorough review prior to proceeding.

Project Organization:

The NLCTA is a project currently being run by the NLC project in the Technical Division.  Dave Burke is the NLC Project Leader and NLCTA Facility Head, Marc Ross is the Operations Manager, and Keith Jobe serves as Operations Manager in Ross’ absence.

Safety Organization:

NLCTA safety organization has Keith Jobe as Safety Officer for almost all aspects of operational safety and project planning.  

· The NLCTA Safety Officer is involved with the design and specification of the 8-pack development.  

· No NLCTA safety oversight of the ORION project design and development has been established to date.

The Accelerator Department Safety Office (ADSO), with Mike Saleski and Paul Miller, is responsible for the operational radiological safety issues such as approval of Radiation Safety Work Control Forms (RSWCF) and control of the ADSO locks on certain radiation devices such as the electron gun pulser and devices requiring control through the RSWCF process.  Their help in this area helps avoid the conflict of interest that can occur because the safety officer is heavily involved in the operations program.

End Station B is a Research Division building.  The Building Manager is Sandy Pierson who has held the position since late September 2001.  We are working with Sandy to understand the appropriate division of responsibilities.

ESH Participation in the program

Radiation Safety issues and authorizations are handled by the Radiation Physics department.  Ralph Nelson is the current Radiation Physicist, with Sayed Rokni serving as both the SLAC Radiation Safety Officer and backup Radiation Physicist.  Support from the OHP group is excellent.

For non-radiological issues and support, the NLCTA receives help from the ES&H division only upon request.  We do not have an ES&H representative who is familiar with the activities or regular operation of the NLCTA, or anyone who is in regular contact with us regarding ongoing issues.

Proposed facility developments:

Near term accelerator beamline development:

· Development of a intra-beam feedback scheme applicable to linear collider final focus beamlines.  (NLC)

· Powering accelerator structures with power derived from the 8-pack modulator installation (NLC)

Near term accelerator support environment development:

· 8-pack installation in north half of end station hall (oil, seismic, electrical, occupational, radiation) (NLC)

· Installation of new generation SLED-II delay line (NLC)

· Installation of delay-line distribution system (NLC)

Components of the NLCTA

The accelerator - beamline

· A shielded accelerator area (“tunnel”) that starts inside End Station B and extends substantially east outside the end station.

· A single PPS controlled area with two entry modules.

· No beamline entering or exiting the tunnel.

· A thermal electron gun without any subharmonic (or other types of) buncher,

· A 60 MeV injector powered by a single X-band klystron,

· A bunch-length defining collimating chicane (at 60MeV)

· Six experimental structure test locations (each able to accommodate a single 1.8 meter structure).  The existing installation uses 3-4 of the test locations for high power structure testing.

· An energy spectrometer (horizontal), with

· A vertical kicker which can spread the beam vertically onto the energy screen to establish correlations between microbunch location, energy, and energy spread (plus some perhaps higher order terms such as distribution shapes).

The accelerator – support equipment

· A thermal cathode DC gun.  This gun runs at about 150 KV and may be driven with either a 1.5 nS or a 150 nS cathode modulator board.  Hardware changes are required to change pulse length.  Peak beam current is about 6 amps with a three-fold hardware, administrative, and BAS limit of allowed repetition rate to 10 Hz or less.

· Three klystron modulators powering differing configurations of pulsed X-band (11.424 GHz, or 4 times SLAC frequency) klystrons.  The first modulator (Station 0) powers a single klystron (30-40 MW pulsed, 1.5 uS @ 60 Hz) which powers the 60 MeV injector.  The second and third modulators (Stations 1 and 2) each have a pair of klystrons which are summed together to generate peak powers of around 100 MW (pulsed, 1.5 uS at 60 Hz).

· Three SLED-II resonate delay line rf compressors.  These compress the incoming rf drive by a factor of 4.5 in rf power, transforming a 1.5 uS pulse to a 250 nS pulse.

· Control racks, control room, etc.

· Magnet power supplies for klystron and accelerator beamline magnets (electrical, seismic)

· Klystron modulator power supplies (electrical, seismic, radiation)

Operation of the NLCTA

Procedures and Operations Documents

NLCTA Operations documents are available on the NLCTA website
.  These documents fall into four general classes:

· Administrative – directives, personnel plan, organization chart.

· Safety – procedures and processes directly related to the safe operation of the facility.

· Training – for operators of the facility.

· Emergency Procedures – the institutional facility emergency plan plus some NLCTA specifics.

The goal is to have the latest approved version of every document online.  Just noted are some exceptions which may be resolved prior to the review. 

A few documents require further discussion:

Operations Directives

This document was modified as part of the change of operation related to unattended operation.  Currently signed off, the document requires some minor changes prior to being released.  (Revision number, org. chart, minor changes).  The most recent version is available in the NLCTA control room.

Beam Authorization Sheet (BAS)

This master document is provided by Radiation Physics and specifically allows the operation of the accelerator as a radiation source at SLAC.  The BAS must be understood by every operator, and initialed on a daily basis to indicate the “safety envelope” conditions are understood and remain valid
.

Daily Inspection Checklist

The conditions for unattended operation (operator not required to be present) with rf (and no beam) require the completion of a checklist every weekday.  The project just received authorization to run over weekend days and vacation days without completing this checklist over the weekends.  This checklist asks that certain radiation shielding controls and shielding be verified.

Weekly Inspection Checklist

The weekly inspection checklist asks that environmental safety, radiation safety, and certain operations conditions be verified.  Since the end station has secondary containment and hazardous waste storage areas, some of the requested checks are required on a periodic basis.

Training and Operator Control

The procedures include sections detailing training for operators of the facility.  All current NLCTA operators have come from MCC either as operators or as accelerator physicists / support personnel, allowing for an easy training experience.  The training program will be revisited as we train the first post-doc operator candidate this coming year.

Operators are issued a key to the NLCTA keysafe.  The keys to the PPS panel are kept in the keysafe, and are only available to NCLTA operators.

NLCTA Safety Systems

The instrumented safety systems for the NLCTA are largely limited to the control of radiological and electrical hazards, fire, and a small system designed to help control a possible environmental spill.

· PPS and Tunnel electrical safety systems follow the best practices of the PPS systems at SLAC.

· A small number of rack mounted power supplies are protected using interlocked doors to limit potential exposure to the rear power supply connections.

· The klystron modulators and 500 kW power supplies have door interlocks and shorting contactors to deactivate and discharge the 20 and 40 kV power sources.

· The fire alarms are signaled to the SLAC fire system.  Additionally, alarms from the high sensitivity smoke detector for the modulators and rf racks turns off the power to the modulators and high voltage power supply.  This interlock was added to reduce the possibility of an arc-driven capacitor fire.

· A level indicator in the secondary containment tanks of the oil tanks of the modulators alerts surveillance of excessive fluid in the tank.  The hazard is not oil spills (for which the containment meets code) but water leaks which have the potential of washing out the entire oil supply with site LCW.

Radiation Safety Systems

The test accelerator is equipped with a relatively simple Personnel Protection System (PPS) and Beam Containment System (BCS) combined with an (electrical) hazard control system.  

The PPS and BCS systems are certified every six months.  The most recently filled out certification procedure is available in the control room.  Additionally, limited PPS checks are preformed following every significant access period, and BCS checks are required every shift when the electron gun is operated.

PPS

Since the NLCTA consists of a single radiological enclosure without any beamlines entering or exiting the enclosure, the PPS system consists of a relatively simple perimeter control system with two fully instrumented PPS gates (“entry modules”).  This is a stand-alone system with no conventional beam stoppers or integration with any other accelerator protection system.

There are four access states:

· Permitted Access – all electrical hazards (magnets) off, gun off, high voltage power supplies for modulators off.  Doors are unlatched, and search secure status is lost.

· Controlled Access – all above listed hazards are off, and must be off to enter this state.  Door is latched and under operator and keybank control.  Area may be searched, setting search secure status.  

· Restricted Access – can only be entered from Controlled Access if search secure is set, and can only be entered from No Access if hazards (electrical and radiation) are off.  In this mode, there are provisions to energize electrical hazards (magnets) with people in the tunnel under special administrative and key controls.

· No Access – accessible from Restricted Access if keybanks are complete (i.e.: tunnel is empty).  Electrical hazards may be turned on.  Radiation hazards require an additional key enable, and may be turned on / remain on only if the Beam Shutoff Ion Chambers (BSOIC) are clear. 

The BSOICs, required to permit the generation of radiation, are set by Radiation Physics to levels described in the Beam Authorization Sheet (BAS).  Operations may monitor and reset the BSOICs, but cannot adjust the detector’s sensitivities.

BCS Safety System

The BCS is designed to disable the high voltage power supplies for the klystron modulators when the beam loss inside the tunnel exceeds the design average beam loss power levels.  The BCS consists of ten ion chambers with approximately ½ liter gas volume installed inside the radiation enclosure.  Trip thresholds are set to 10 nA.

Operations documents describe the daily checks required for beam running.

Fire Safety System

The NLCTA is located inside End Station B.  This building is equipped with rated ceiling-mounted ionization smoke detectors which may be able to provide early detection of an equipment fire.

Parts of the NLCTA are sprinklered.  These parts include the radiation enclosure, the power supply and rf racks, and the control room.  The sprinklers are not expected to respond in time to significantly limit equipment damage in the event of a fire.  Because the entire end station is not sprinklered, the building is not considered as a sprinklered building per code.

The project is equipped with a high sensitivity smoke detection (HSSD) system which is installed on most of the racks supporting the klystrons and low level rf electronics.  Coverage also includes the high voltage power supplies for the klystron modulators, klystron modulators, and the radiation enclosure.

In addition to the SLAC Fire Alarm System, one HSSD channel (1) removes interlock permissive for the modulator power supply and (2) trips the 208 volt circuit breakers supplying modulator power.

General Safety Concerns and Areas of Interest

Environmental Safety & Hazardous Materials

The project uses 2000 gallons of mineral oil as a dielectric insulator in the pulse tanks for the three modulators.  The development of the 8-pack will double the oil use. Mineral oil is a NFPA Hazard Identification: Health 0, Fire 1, Reactivity 0 and is rated as a CA Hazardous Material.

Procedural measures:

Spill response includes a 55 gallon wet-dry vacuum cleaner designated for emergency use and several hundred pounds of absorbent.

Engineering controls:

Tunnel access has been modified to prevent spilt oil from running down into the subterranean tunnels. 

Active protection equipment:

Level sensors in the secondary containment for the modulators signal in the Sector 30 gate.  Procedures indicate responses necessary to secure the water supply to prevent a water leak from washing oil from the primary containment, over the secondary containment, and into the environment.

Other than oil, there are no NLCTA-specific environmental issues that are different than commonly encountered throughout the institution. 

Pressure and Vacuum Vessel

Not applicable.

NLCTA does not have any pressure or vacuum vessels that require review by the Pressure and Vacuum Vessel Safety citizen’s committee
.  The SLED delay lines and the vacuum transport lines are classified as distributed systems and are thus exempt.

Laser Safety

There are no high power lasers in the NLCTA at present.  

Small He-Ne and diode laboratory lasers are sometimes used for alignment and calibration of diagnostic equipment.

Seismic Safety

The following presentations have been made to the Seismic Safety Committee with approvals granted:

8/27/92
NLCTA Shielding Enclosure
5/27/93
End Station B South Wall Reinforcement
8/11/93
B501 Power Substation
9/13/93
NLCTA Overhead Utilities Cable Tray System
5/11/94
NLCTA Control Building

There is no evidence that the support equipment, power supplies, racks, or experimental equipment have ever been reviewed or inspected by the committee.

The NLCTA items of seismic interest are:

· Racks are reasonably secured.  Connection to the concrete deck appears to have been properly made using Hilte Quickset fasteners.

· Power supplies (both LGPS type and 500 KW type) are restrained in a “SLAC-Standard” manner with “dogs” (plates with a bolt at the fulcrum).  While these provide some measure of seismic restraint, they do not meeting the committee’s current standard.  Since this is an institutional problem and our power supplies are interchangeable with others on site, we have chosen to have this to be addressed in a site-wide and not a project-specific manner.  At this point, there is no site-wide solution for securing the this class of power supplies.

· Accelerator equipment and rf transport systems are believed to be reasonably restrained.

· Klystrons are not restrained (they are well supported against gravity).

· There is no information on the seismic competency of the bridge crane.

· The rolling door has not been seismically evaluated.  There is a small chance that the top-side clearance is small enough to ensure that the wheels will not jump the rails.  (See next item).

· The East “Plug Wall,” North, South, and parts of the East walls of the End Station probably fail seismic standards in a significant manner.  Additionally, the North Emergency Exit (the secondary life safety exit for the building) will certainly collapse given a significant tremor.  A “Phase II Seismic Assessment” of End Station B is being considered by Brad Youngman.

Crane Safety:

The End Station has the original two-hook 15 / 50 ton bridge crane equipped with a relatively modern remote radio controller.  The vast majority of the crane use is for devices which are a very small fraction of the hook’s rated capacity.  The situation is as follows:

· There is no institutional record of a certification ever having been done on the crane
.  The installation company almost certainly certified the crane at the start of operation in the mid 1960’s, but there are no records to support this statement.  The DOE recertification period for this type of crane is believed to be three years.

· The bridles in use have been recently certified and dated by an outside contractor.

· The maintenance of the crane has not been reviewed or carefully organized.  The SEM maintenance personnel respond quickly and thoroughly to specific reports of problems and regularly “service” the crane, but do not have a strong maintenance program or one supported by a licensed inspector.  Frank Brenkus has just posted a inspector position on site, and we hope the situation will improve.

Radiation Safety

The NLCTA was designed and approved based on a higher beam energy, current, and power than is currently deliverable.  Radiation Physics reviews the operation of the accelerator every 6 months as part of the renewal of the Beam Authorization Sheet.  Ralph Nelson of Radiation Physics will address the committee on the status of the radiation safety program.

Electrical Safety, AC

The AC distribution system has not undergone any significant changes since the facility was initially commissioned.  The power backbone was initially installed with plans for future changes in program scope and size, and has served the facility well.

Grounding

The NLCTA is a well bonded system with verifiable bonding on all accelerator instrumentation, components, power supply, and support systems.

Multiple power sources

The klystron modulator high voltage power supplies and more specifically the modulator itself are examples of devices with multiple power sources requiring the use of a ELP.  The latter is powered separately by (1) a single conductor 25 amp 20 kV circuit with stored energy in excess of 10 kJ and (2) a 50 amp 3 phase 208 volt power source. 

Bakeout

The NLCTA is the only facility at SLAC which regularly uses equipment ground fault detectors and interlocks on 110 and 208 volt powered bakeout systems.  Experience has shown that this both protects equipment from uneven heating or burn-through, but also provides a significant reduction in electrical shock potential.

Electrical Safety, DC

Three classes of DC electrical safety issues exist in the NLCTA.

· Behind the rack distribution.  All power supply connectors are either covered, behind interlocked doors, or at legal voltage / current.  

· Tunnel hazards.  All exposed tunnel electrical hazards (there are 100 uncovered magnets and 1 gun) are interlocked to the PPS Hazard Control program which is regularly certified.  Additionally, the NLCTA Procedure 02-03-10 details the locking out of all hazards with a dedicated lock set with a single securable key.  This key is locked in the NLCTA control room by all affected personnel with their personal locks and tags during installation periods.

· Modulator power.  The modulator and it’s power supply is interlocked with a number of door interlocks which turn off the hazardous energy source and short the stored energy.  An ELP exists for these devices.

Lock and Tag remains a continuing challenge.  Due to the PPS interlocking, the determination of source to load control is difficult to verify, and thus the procedure takes a broad approach and locks out the entirety of the tunnel hazards.

Occupational Safety

SLAC safety guidelines cover all work in the NLCTA area.  Special hazards in the project which require special attention include fall protection for lighting and crane maintenance.  Groups doing high risk work are expected to be knowledgeable and supply the appropriate personal protection equipment (PPE).

NLCTA Shop

Trailer 485 (a converted shipping container) houses the NLCTA workshop.  Included in the shop are a variety of hand and portable power tools.  Additionally, the shop contains a drill press, sander, and an oxygen - acetylene torch.  Outside the shop are a variety of cutoff tools which, while used by NLCTA, are part of the Research Division’s tool assortment.

Users of the shop are primarily physics and engineering support people associated with the project.  On the job training covers all equipment use.

NLCTA Storage Buildings

Trailers 483, 484, and 486 are shipping containers used to store cable, power supply components, and vacuum components respectively.  There are no identified safety issues in the storage buildings.

ADA Accessibility

The control room is minimally accessible per the current standards for persons with mobility disabilities.  The experimental area, beamline, racks, and support environment are not accessible per the ADA criteria.

� � HYPERLINK "http://www-project.slac.stanford.edu/lc/local/Projects/8Pack/8pack.html" ��http://www-project.slac.stanford.edu/lc/local/Projects/8Pack/8pack.html�


� � HYPERLINK "http://www-project.slac.stanford.edu/orion/" ��http://www-project.slac.stanford.edu/orion/�


� http://www-project.slac.stanford.edu/lc/local/Projects/NLCTA/nlcta.htm


� The concept of “daily” signatures on the BAS is currently being redefined to allow over-weekend unattended operation without beam.


� Richard Boyce, private conversation.  10/29/2001


� Liam Robinson, SEM.  Private conversation, 10/2001.





Z:\nlcta\NLCTA Safety Audit 1.1.doc
10/31/2001 21:25
page 3 of 14

